NAUDOJIMO INSTRUKCIJA	mALB

VITROS mALB Reagentas	Mikroalbuminas

680 1740

Paskirtis

Tik in vitro diagnostikai.

VITROS mALB Reagentas skirtas VITROS 5,1 FS sistemai ir VITROS 5600 Integruotajai Sistemai, kiekybiniam albumino koncentracijos žmogaus šlapime (mALB). Šlapimo albumino matavimas padeda diagnozuoti diabetinę nefropatiją, hipertenziją ir kardiovaskuliarinį sutrikimą.

Tyrimo apžvalga ir paaiškinimas

Mikroalbuminurija yra charakterizuojama pagal šlapimo albumino išskyrimą virš normalių lygių1, 2, 3 Vidutiniškai padidėjęs albumino išskyrimas buvo pripažintas prognozuoti apie artėjančią inkstų ligą pacientams turintiems hipertenziją arba diabetą.4  Mikroalbuminurija yra pastebima, jeigu šlapimo albuminas yra ribose 30–300 mg/ 24 h.5

Procedūros principai

Kiekybinis šlapimo albumino matavimas yra atliekamas naudojant VITROS mALB Reagentą kartu su VITROS Kalibratoriumi “Kit 24” sistemose VITROS 5,1 FS ir VITROS 5600 Integruotoje. VITROS mALB Reagentas yra dvigubos kameros pakuotėje, kurioje yra stabilūs, paruošti naudojimui skysti reagentai. Mėginiai, kalibratoriai ir kontrolės yra sumaišomos su Reagentu 1, turinčiu polimerą ir paviršinę medžiagą. Antiserumas žmogaus albuminui (Reagentas 2) iššaukia imunocheminę reakciją duodančią antikūnio/antigeno kompleksus. Šviesos sklaidos savybės keičiasi proporcingai nuo antikūnių/antigenų kompleksų, padidinamas tirpalo drumstumas, priklausomai nuo albumino koncentracijos mėginyje. Drumstumas matuojamas spektrofotometriškai 340 nm bangos ilgiu. Kai kalibracija jau atlikta kiekvienai reagento partijai, šlapimo albumino koncentracija kiekviename nežinomame mėginyje nustatoma naudojant išsaugotą kalibravimo kreivę.

Testo Tipas ir Sąlygos

Testo Tipas	VITROS Sistema

Apytikris

Inkubacijos Laikas	Temperatūra	Bangos ilgis

Inkubacija 1: 5

Reakcijos Mėginio

Tūris

Vertinimo kriterijus	5600, 5,1 FS

minutės	37 °C (98.6 °F)	340 nm	12 µL Inkubacja 2: 5

minutės

Įspėjimai ir atsargumo priemonės

Tik in vitro diagnostikai.

ĮSPĖJIMAS:	Su mėginiu, skystomis ir kietomis atliekomis elkitės kaip su potencialiai infekuotomis, pagal CLSI nurodymus M29 ar kitą išleistą Biologinio pavojaus saugumo informaciją.

Kita specifinė informacija, dėl kalibratorių, kontrolių ir kitų komponentų yra pateikta naudojimo instrukcijose

References

1.	Mogensen, C.E., et. al., “Early Changes in Kidney Function, Blood Pressure and the Stages in Diabetic Nephropathy,”

in Prevention and Treatment of Diabetic Nephropathy, H. Keen and M. Legrain, ed. Boston, MTP Press, pp. 57-83, 99-105, (1983).

2.	Mogensen, C.E., C.K. Christensen and E. Vittinghus, “The Stages in Diabetic Renal Disease, with Emphasis on the Stage of Incipient Diabetic Nephropathy,” Diabetes, 32:Suppl. 2, 64-78, (1983).

3.	Viberti, G.C., et. al., “Proteinuria in Diabetes Mellitus: Role of Spontaneous and Experimental Variation of Glycemia,”

Kidney International, 21:714-20, (1982).

4.	Burtis C, Ashwood, ER (ed). Tietz Textbook of Clinical Chemistry, 3rd ed. Philadelphia, PA; WB Saunders Co; 484; 1999.

5.	American Diabetes Association. “Standards of Medical Care for Patients with Diabetes Mellitus” Diabetes Care, 23:Suppl.

1, S32-42, (2000).

6.	CLSI. Protection of Laboratory Workers from Occupationally Acquired Infections; Approved Guideline – Third Edition.

CLSI document M29-A3 (ISBN 1-56238-567-4). CLSI, 940 West Valley Road, Suite 1400, Wayne, PA 19087-1898 USA;

2005.

7.	Calam RR. Specimen Processing Separator Gels: An Update. J Clin Immunoassay. 11:86–90; 1988.

8.	NCCLS. Urinalysis and Collection, Transportation, and Preservation of Urine Specimens; Approved Guideline – Second

Edition. NCCLS document GP16-A2 (ISBN 1-56238-448-1). CLSI, 940 West Valley Road, Suite 1400, Wayne, PA

19087-1898 USA; 2001.

9.	Manley, et al. Decreases in Albumin/Creatinine and N-Acetylglucosaminidase/Creatinine Ratios in Urine Samples Stored at -20 °C, Clin Chem, 38/11, 2294–2299; 1992.

10.    European Commission.The Certification of a Matrix Reference Material for Immunochemical Measurement of 15 Serum

Proteins, ERM-DA470. Report EUR 15243 EN and 16882 EN, European Communities, 2004.

11.    CLSI. Statistical Quality Control for Quantitative Measurements: Principles and Definitions; Approved Guideline – Third

Edition. CLSI document C24-A3 (ISBN 1-56238-613-1). CLSI, 940 West Valley Road, Suite 1400, Wayne, PA 19087-1898

USA; 2006.

12.    NCCLS. Interference Testing In Clinical Chemistry; Proposed Guideline. NCCLS document EP7-A [ISBN 1-56238-480-5].

CLSI, 940 West Valley Road, Suite 1400, Wayne, PA 19087-1898 USA; 2002.

13.    Young DS. Effects of Drugs on Clinical Laboratory Tests. Ed. 4 Washington, D.C.: AACC Press; 2000.

14.    Friedman RB, Young DS. Effects of Disease on Clinical Laboratory Tests. Washington, D.C.; AACC Press; 1997.

15.    NCCLS. Method Comparison and Bias Estimation Using Patient Samples; Approved Guideline. NCCLS document EP9- A2 [ISBN 1-56238-472-4]. CLSI, 940 West Valley Road, Suite 1400, Wayne, Pennsylvania 19087-1898 USA; 2002.

16.    NCCLS. Evaluation of Precision Performance of Quantitative Measurement Methods; Approved Guideline - Second

Edition. NCCLS Document EP5-A (ISBN 1-56238-368-X). CLSI, 940 West Valley Road, Suite 1400, Wayne PA, USA,

19087-1898, 2004.

17.    NCCLS. Evaluation of Precision Performance of Quantitative Measurement Methods; Approved Guideline – Second

Edition. NCCLS document EP5-A2 [ISBN 1-56238-542-9]. CLSI, 940 West Valley Road, Suite 1400, Wayne, PA

19087-1898 USA, 2004.

Ortho-Clinical Diagnostics

Johnson & Johnson

50-100 Holmers Farm Way High Wycombe Buckinghamshire

HP12 4DP United Kingdom

Ortho-Clinical Diagnostics, Inc.

100 Indigo Creek Drive

Rochester, NY 14626-5101

Brasil:

Johnson & Johnson do Brasil Ind. e Com. de Prod. para Saúde Ltda.

Divisão: Johnson & Johnson Medical Brasil

Rod. Presidente Dutra, Km 154, São José dos Campos - SP - CEP: 12240-908 - Brasil

CNPJ: 54.516.661/0002-84

Resp. Técnico.: Nancy M.R.B Lopes C.R.F.-SP Nº 10965

Somente Para Uso Diagnóstico In Vitro

SAC: 08007075420

VITROS is a registered trademark of Ortho-Clinical Diagnostics, Inc.

© Ortho-Clinical Diagnostics, Inc., 2004-2009.