[image: image47.png]




Microchip - precision pressure catheter with semiconductor pressure sensors for reliable epidural, ventricular, and parenchymal 
intracranial pressure measurement

Instruction manual

for neurosurgical precision pressure catheters

with and without

additional temperature and

oxygen measurement
(multi-parameter catheter)
Table of contents

21. General description


22. Indications


23. Counter indications


34. Side effects


45. Overview of RAUMEDIC® precision pressure catheters


56. Precautionary measures / notes


77.1. Correct connection and removal


77.2. Zeroing


78. Making a drill hole


89. Implantation with the BOLT KIT


89.1. Catheter type specific notes on implantation


910. Implantation with spliceable tunnelling sleeve


1010.1. Catheter type specific notes on implantation


1111. Notes on the implantation of NEURODUR®, -TEMP


1112. Notes on the implantation of NEUROVENT®-VP16


1213. Fixing the catheter in place after implantation


1214. Safety information


1315. MRI safety RAUMEDIC® Precision pressure catheters at 1.5 T / 3.0 T


1415.1. Use during MRI


1515.2. MRI parameters


1616. Indication and warning symbols on the label


1617. Technical data


1818. Accessories


1919. Liability and guarantee information


1920. Disposal




1. General description
The neurosurgical pressure catheters by RAUMEDIC® are Class III medical devices according to EC Directive 93/42/EEC that may remain in the body for a maximum of 29 days, even with cumulative application. 

The neurosurgical pressure catheters are precision measuring instruments with integrated semiconductor pressure sensors. They are used to make high-resolution and artefact-free images of intracranial pressure curves during temporary or continuous pressure monitoring up to frequency analysis.
In addition, RAUMEDIC® pressure catheters come with an optional integrated thermistor for the additional measurement of the local brain temperature at the pressure measurement site.

Another parameter that can be additionally measured with the appropriate RAUMEDIC® pressure catheter is the oxygen partial pressure (ptiO2) of the interstitial fluid. The ptiO2 permits conclusions about the prognosis of secondary brain damage. A highly accurate optical measurement method (quenching of luminescence) is used for that.

2. Indications
The measurement of intracranial pressure (ICP) and cerebral perfusion pressure (mean arterial blood pressure minus ICP) are, according to the current state of knowledge, essential requirements for effective treatment of patients with suspected intracranial pressure increases (e.g. craniocerebral injury GCS ( 8; malignant media infarction; hepatic encephalopathy, subarachnoid hemorrhage Hunt/Hess grade IV + V; cerebral oedema; hydrocephalus) or other patients amongst whom the clinical picture may be associated with an increase in the ICP.

The measurement of oxygen partial pressure ptiO2 enables additional information about the patient's condition (secondary brain damage – detection of cerebral hypoxia and tissue damage). According to the general opinion in specialised literature, a good oxygen supply to the brain tissue correlates with a good outcome for patients.

The measurement of the brain temperature allows the direct measurement of the temperature in the brain tissue. This temperature can differ from the core body temperature. Measurement of the brain temperature allows the detection of episodes of cerebral hypothermia and the derivation of additional temperature information with hypothermia treatment.
3. Counter indications

The following notes are no replacement for studying experience reports and recommendations in specialised medical literature that deal with techniques in pressure catheter application, failure and malfunctions as well as complications that occur.

There must be a particularly strong indication for the use of RAUMEDIC® precision pressure catheters with:

· Reye’s syndrome

SYMBOL 183 \f "Symbol" \s 10 \h
Clotting disorder (coagulopathies) due to an increased risk of haemorrhages

· Application with babies and toddlers < 1 year
· Sepsis, infection (especially in the catheter entry point area)

· Encephalitis

· Lysis therapy (haemorrhaging risk)

4. Side effects

Infections, liquor fistulas and neurological after-effects are potential complications of the process. Bone, meningeal, cortical or intraparenchymal haemorrhages can also occur around the implant site.
5. Overview of RAUMEDIC® precision pressure catheters

	Catheter description
	Measuring site
	Functions
	Application technique

	Catheter designation
	Art.No.
	Ø (CH)
	Parenchyma
	Ventricle
	Epidural
	ICP
	TEMP
	ptiO2
	Drainage
	Bolt
	Tunnel
	Auxiliary devices

	NEUROVENT®-P
	092946-001
	5 F
	
	
	
	
	
	
	
	
	
	

	NEUROVENT®-P-TEMP
	094268-001
	5 F
	
	
	
	
	
	
	
	
	
	

	NEUROVENT®-PTO
	095008-001
	5 F
	
	
	
	
	
	
	
	
	
	

	NEUROVENT®-PTO 2L
	095108-001
	5 F
	
	
	
	
	
	
	
	
	
	

	NEUROVENT®-TO
	095908-001
	3 F
	
	
	
	
	
	
	
	
	
	

	NEUROVENT® with sleeve housing
	091576-001
	9 F
	
	
	
	
	
	
	
	
	
	mandrel (soft)

	NEUROVENT®-VP16
	096704-001
	9 F
	
	
	
	
	
	
	
	see section 12.
	stylet

	NEUROVENT®
	092956-001
	9 F
	
	
	
	
	
	
	
	
	
	stylet

	NEUROVENT®-TEMP
	094278-001
	9 F
	
	
	
	
	
	
	
	
	
	stylet

	NEUROVENT®-IFD-S
	091678-001
	9 F
	
	
	
	
	
	
	
	
	
	mandrel (soft)

	NEUROVENT®-TEMP-IFD-S
	094288-001
	9 F
	
	
	
	
	
	
	
	
	
	mandrel (soft)

	NEUROVENT-IFD-R
	095317-001
	9 F
	
	
	
	
	
	
	
	
	
	mandrel (hard)

	NEUROVENT®-TEMP-IFD-R
	095327-001
	9 F
	
	
	
	
	
	
	
	
	
	mandrel (hard)

	NEUROVENT® 6F
	094678-001
	6 F
	
	
	
	
	
	
	
	
	
	mandrel (soft)

	NEURODUR®
	092976-001
	-
	
	
	
	
	
	
	
	see section 11.
	

	NEURODUR®-TEMP
	094298-001
	-
	
	
	
	
	
	
	
	see section 11.
	


6. Precautionary measures / notes

· Neurosurgical precision pressure catheters are precision measurement instruments and may be used by qualified personnel only!

· The neurosurgical precision pressure catheters are intended for single use only.
If reused, there is a risk of prion transfer (Creutzfeldt-Jacob disease) as well as damage to the catheter from reuse processes.
· An optical check of packaging and of the pressure catheter for mechanical integrity as well as a check of the date of expiry must be carried out before use! Do not use the pressure catheter if checking indicates damage or evidence of incorrect handling! Use is also prohibited if the date of expiry is exceeded!
· The catheter is sterile and must be implanted under sterile conditions in the operation room!

· Prior and during the examination, the pressure catheter must not be bent, stretched or squeezed as well as manipulated by surgical tools. This is particularly true of catheters with an integrated partial oxygen pressure metering function. ( Caution: risk of glass fibre breakage!

· If an unchanged ICP is indicated given the absence of an intracranial pressure pulse amplitude (breath-synchronously variabilities / pulse peaks) in a vigorous patient, the functionality of the pressure catheter has to be checked immediately by testing the reaction of the pressure channel on a slight movement of the head, for example. If the signal does not change, the defective precision pressure catheter has to be replaced.
[image: image48.wmf]
· The pressure sensors may not be subjected to undefined pressures, e.g. by gripping the catheter with forceps to insert into a hole, because the sensors may be damaged with pressure values in excess of 1500 mmHg!

Do not touch the membrane and do not exert pressure with fingers!
· The catheter entry point has to be checked for any signs of a wound infection in regular intervals (such as swelling, reddening). If there are any signs of an infection, the catheter has to be removed.

· Since the infection rate of up to 3 days for a catheterisation period is very low, it is generally recommended to change a catheter after 5 days and no later than 10 days (see Citation 1, last page below).
· It must be insured with implantation that the sensor can be removed without problems with a slight pull.

· Extreme pull or damage to the catheter with sharp instruments has to be avoided – in particular during removal. Check for completeness after having removed the catheter.

· Bone, meningeal, cortical or intraparenchymal haemorrhages can occur around the operation area. If there is a suspicion of such occurrences, an imaging check-up procedure should be carried out!
· A therapy decision, based on the oxygen partial pressure value alone must not be made, rather it should always be backed up with another method, e.g. ICP measurement.
7. Precision pressure catheter connection to monitor systems

Pressure measurement:

[image: image49.jpg]


The measurement chain for pressure measurement comprises:

[image: image50.png]



- RAUMEDIC® precision pressure catheter (1)

- ICP-TEMP-Cable (2)

- Zero point module NPS2 (3)
[image: image1.png]


 Note:
The zero point module NPS2 must be selected to match the configuration of the existing monitor system in the clinic.


( see 18. Accessories

Temperature measurement:
The measurement chain for additional temperature measurement comprises:

[image: image51.png]


- RAUMEDIC® precision pressure catheter (1)

- ICP-TEMP-Cable (2)

- ICP-TEMP-Adapter (4)

- Zero point module NPS2 (3)
For additional temperature measurement, an ICP‑TEMP-Adapter is inserted between ICP-TEMP-Cable and NPS2.

[image: image2.png]


 Note:
The ICP-TEMP-Adapter must be selected to match the configuration of the existing monitor system in the clinic.

[image: image52.png]Preccrmsre e,
Sty
-
-

" -

2 cm Dura

fixation }

detached
: probe
T ramp(/

ligature
groove




( see 18. Accessories

Oxygen partial pressure measurement:

The measurement chain for oxygen partial pressure measurement comprises:

- RAUMEDIC® precision pressure catheter (1)

- Cable PTO
- Cable LWL (Optical fibre cable)
- DATALOGGER MPR2 logO / EASY logO

For additional oxygen partial pressure measurement, ICP-TEMP-Cable and Adapter are replaced with a Cable PTO. In addition, a LWL Cable (Optical fibre cable) is required for connection of the oxygen measurement channel. 

[image: image53.png]right



[image: image3.png]


 Note:
The oxygen can be measured only with RAUMEDIC® Datalogger MPR2 logO or EASY logO. A NPS2 zero point module is no longer required


( see 18. Accessories

7.1. Correct connection and removal

The gold dots on the plug connections of ICP-TEMP or Cable PTO and catheter must point to the same side.

[image: image54.png]protective cap

Cle

D





          [image: image4.png]gold dots
vy
b 7
N




Joining and releasing the plug connection

[image: image55.png]



When the optical plug connection comes into contact via LWL Cable (Optical fibre cable), the lateral nose on the male OF connector must be pushed into the appropriate recess of the counterpart all the way to the stop (1). Subsequently, the rim must be rotated clockwise 90° until it latches (2) to secure the plug connection. The plug connection is released in reverse order. 
[image: image56.png]



[image: image5.png]


 Caution:
To avoid incorrect measurements, the optical fibre plug connection (OF contact) for ptiO2 measurement must be checked to ensure that is free from dirt!
7.2. Zeroing
Zeroing must be done according to the NPS2 zero point module instruction manual (zero offset adjustment of the monitor while pressing the blue button of the NPS2).
Oxygen partial pressure measurement via DATALOGGER MPR2 logO or EASY logO does not require zero offset adjustment. All necessary calibration data for correct display of values are stored in the catheter connector and are automatically accepted by the MPR2 logO after the PTO connection cable is connected.

[image: image6.png]


 Note:
After catheter removal, a possible catheter drift during the recent implantation period can be detected by reading out the pressure value under zeroing conditions (immersion in fluid, no direct light).
8. Making a drill hole

Holes must be drilled into the skullcap according to the BOLT KIT / DRILL KIT instruction manual.

[image: image7.png]


 Caution:
The operating area must be freed from bone splinters or treatment residue after every drilling process or any work on the drill hole (e.g. shaping a ramp) with rinsing. Possibly existing burrs at the base of the drill hole must be removed carefully. 
9. Implantation with the BOLT KIT

[image: image8.png]


 Caution:
Please refer to the RAUMEDIC®- precision pressure catheter overview (see section 5.) to find out if the catheter that is used is compatible with the bolt system.

The Bolt KIT version must be chosen according to the catheter to be used:

	BOLT KIT CH5
	BOLT KIT CH9
	BOLT KIT PTO
	BOLT KIT PTO 2L

	for all catheters with a diameter of 5F and without ptiO2 measurement function
	for all catheters with a diameter larger than 5F
	for

NEUROVENT®-PTO,

NEUROVENT®-TO
	for

NEUROVENT®-PTO 2L


RAUMEDIC®- precision pressure catheter application with BOLT KIT must be done according to the BOLT KIT / DRILL KIT instruction manual. Additional notes to be observed during application depending on catheter type are described in the following sections.

The BOLT KIT fixing cap must be turned until the catheter is secured and a complete seal is made. It must also be ensured that the catheter remains open for ventilation and drainage. To safely determine this position with a laid down mandrel, the crimp connection must be closed until the mandrel can still barely be pulled out.

9.1. Catheter type specific notes on implantation
[image: image57.png]


NEUROVENT®-PTO, -TO and -PTO 2L

These catheter models have a mechanical stop (1) in the front region of the catheter. This serves to place the catheter tip in ideal implantation depth (2.5‑3.0 cm) if BOLT KIT PTO or BOLT KIT PTO 2L is used:

The lip on the catheter stopper must be pushed forward by the silicone ring (sealing cap not tightened). During passage of the silicone ring, a resistance must be felt against which the catheter must be pushed forward until the silicone ring has completely passed through and the base of the Bolt screw has been reached.
The following image illustrates the result (example illustration of the NEUROVENT®-PTO in the BOLT KIT PTO):
[image: image58.wmf]There is a similar image for the NEUROVENT®-PTO 2L in the BOLT KIT PTO 2L, however this catheter does not have a lip on the catheter stopper in order to allow tunnelled application. The catheter stopper must be pushed forward through the silicone seal up to the base of the Tuohy Borst adapter of the BOLT KIT PTO 2L.
The catheter is subsequently fixed in place in the BOLT screw / Tuohy Borst adapter by tightening the fixing cap.
NEUROVENT®-IFD-R and -TEMP-IFD-R

This catheter model may be applied only with the BOLT KIT CH9. Tunnelling of this catheter is not permitted due to the strong restoring force of the mandrel that is inserted into the catheter! After reaching the desired implantation depth, the mandrel must be pulled out of the catheter lumen slowly and carefully.

[image: image9.png]


 Caution:
These catheters are equipped with an additional lumen for the mandrel. It reaches into the metal tip of the catheter and thus allows precise guidance during application. Since this lumen also transmits the reference pressure to the pressure sensor chip, it may not be closed or come into contact with fluid! 

NEUROVENT®-IFD-S, -TEMP-IFD-S and -6F
[image: image10.png]


 Caution:
These catheters are equipped with a mandrel that is taken through the drainage lumen. It must be ensured during application that the top of the guide wire does not push through a proximal drainage opening. After successful ventricle puncture, the guide wire is pulled back and removed.

10. Implantation with spliceable tunnelling sleeve

[image: image11.png]


 Caution:
Please refer to the RAUMEDIC® precision pressure catheter overview (see section 5.) to see if the catheter that is used is suitable for tunnelling.

The version of spliceable tunnelling sleeve must be chosen according to the catheter to be used:

	Spliceable tunnelling sleeve CH8
	Spliceable tunnelling sleeve CH12

	for all catheters with diameter 5F
	for all catheters with a diameter larger than 5F and the NEUROVENT®-PTO 2L


Application of a RAUMEDIC®- precision pressure catheter with spliceable tunnelling sleeve must be done according to the instruction manual for the spliceable tunnelling sleeve with trocar. Additional notes to be observed during application depending on catheter type are described in the following sections.

After implantation of a catheter via tunnelling, it must be pulled back subcutaneously until it rests below a skin-galea flap and loops cannot form. It is then led from the point of exit at the stab incision to the fixation with a slight curve.

[image: image12.png]


 Caution:
During the explantation of tunnelled catheters, the sutures over the drill hole must first be loosened fully. The catheters can subsequently be pulled out like wound drainage using suitable tools (e.g. forceps) for non-destructive drill hole passage. 

[image: image13.png]


 Caution:
If there is a plan not to open the stitches during removal of the catheter due to clinical requirements, suitable measures such as:

· angled drill hole (similar to Section 10.1.) 

· ramp-shaped deburring of the drill hole margin in the drainage direction of the catheter 

· suitable choice of drill hole diameter (depending on the anatomical conditions), but at least: 

· Ø 10 mm for the 5F catheter and 

· Ø 12 mm for all catheters with a diameter larger than 5F and the NEUROVENT® PTO 2L

must be used to ensure that the catheter can be removed. To prevent the catheter from being torn off in the process, pull it slowly and not jerkily.
10.1. Catheter type specific notes on implantation
NEUROVENT®-PTO 2L

[image: image14.png]


 Note:
NEUROVENT-PTO 2L can also be used for craniotomy patients.

The drill hole for the tunneling of this catheter must be made slanted at an approx. angle of 30°. The upper edge of the bone must be broken in the direction of the leading out of the catheter to avoid sharp edges.

[image: image15.png]


 Caution:
The direction of introduction of the catheter should correspond with the direction of the sloping drill hole (30°). The catheter must be taken into the brain tissue straight through the drill hole. For craniotomy patients, the catheter must also be introduced into the brain tissue at an approx. angle of 30° to the surface of the brain. Bending the thin CH5 catheter tube must be avoided (danger of breakage of the O2 glass fibre!).
NEUROVENT®, -TEMP, -IFD-S, -TEMP-IFD-S, -6F and with sleeve housing
	[image: image16.png]



	For ventricle pressure measurement, these catheters can be placed like external CSF drainage. The LL cap must first be removed for that. Leaking liquor indicates correct positioning of the pressure catheter. An external drainage system (e.g. RAUMEDIC®-Liquor-Drainage-Set) can now be connected or the lumen be closed again with the LL cap. Curved incision with tunnelling of the skin-galea flap prevents that skin incision and drill hole are directly above each other.

After tunnelling, the catheter is pulled out about 30 cm on the drill hole side and inserted into the ventricle through the drill hole. 

The catheter can be stiffened for ventricle puncture with a stylet (1) or the internal mandrel (2):

1)
Carefully insert the stylet in the distal drainage opening and push it in up to the measurement head.

2)
Carefully insert the guide wire (mandrel) through the Luer attachment, push it up to the measurement head and lock it in place with Luer cap. 

[image: image17.png]


 Caution:
Make sure that the stylet or mandrel tip does not push through a proximal drainage opening!

After successful ventricle puncture, stylets are pulled out and removed with a slight rotation around their longitudinal axis while guide wires are pulled out and removed continuously holding the catheter above the drill hole.

	[image: image18.png]Messkopf




	


11. Notes on the implantation of NEURODUR®, -TEMP

Placement of the brain pressure measurement probe between dura and skullcap. The measurement window is pointing toward the dura here. Correct alignment can be detected from the outside only with the catheter scale that faces upward.

The minimum size for the drill hole is 9 mm. The edge of the bone in the intended introduction and extraction direction is bevelled and smoothed off to form a tapered edge using milling hand piece, Luer bone rongeur, sharp curette or by creating a second drill hole down to the cancellous bone directly alongside the actual drill hole. Wax can be used for smoothing and haemostasis.

Detach a dura section of about 25 mm length and 25 mm width from the inside of the bone surface with a suitable instrument (e.g. spring probe) to avoid internal stress on the measurement window.

A lateral stab incision is made a few centimetres from the drill hole. After tunnelling, the probe is carefully pushed through with atraumatic forceps (no sharp-edged instruments!) or with a finger and inserted subgaleal into the drill hole. The measurement head must be inserted at least 5 mm below the edge of the bone between dura and skullcap without using force! 

The pressure can be measured at any point on the skull. The present disease should be used for orientation when selecting the site of implantation. Epidural precision catheter application should be done on the side of the lesion to detect pressure changes early.

[image: image19.png]


 Note:
Meaningful values can only be recorded after an adjustment period of about 2-6 hours.
The dura requires this time to shape to the probe. Applanation quality is achieved if pulse and breath amplitudes are transmitted cleanly with bottom peaks.

After neurosurgical operations with trepanation, the NEURODUR® /-TEMP can be used for postoperative brain pressure monitoring without further surgical effort. The measurement head must be pushed below the trepanation edge as far as possible. The dura must be detached here too so that the measurement head slides between dura and cranial bones easily and there are not measuring errors from internal dura tension.

[image: image20.png]


 Note:
The NEURODUR® /-TEMP can also be laid subdural on the surface of the brain.
The catheter channel must be sealed against leaking liquor with deep sutures!
12. Notes on the implantation of NEUROVENT®-VP16

[image: image21.png]


 Note:
We recommend placing the NEUROVENT®-VP16 with BOLT KIT CH9 (see section 9).

The implantation of NEUROVENT®-VP16 can also be done with a sterile skin incision with a sufficiently large drill hole (> 6.0 mm) as an alternative to the recommended method.

The catheter can hereby be stiffened with the included stylet or another disposable stylet (diameter ≤ 1.1 mm, min. length 160 mm, maybe with the possibility for neuro-navigation) for ventricle puncture.
The stylet is temporarily inserted in the drainage lumen for that. After successful puncture, the skin incision is closed and the stylet removed. The drainage lumen is connected with an external ventricle drainage system. 
[image: image22.png]


 Caution:
To avoid infection, the condition of the wound must be monitored especially closely with this application technique!

13. Fixing the catheter in place after implantation

RAUMEDIC® precision pressure catheters must always be additionally fastened to the scalp with the included fixing wing to avoid manipulation of the position regardless of the implantation method used!


RAUMEDIC® precision pressure catheters with a diameter larger than 5F

These catheters must be additionally fixed in place with two ligatures on the existing ligature grooves to protect against longitudinal displacement or rotation!
NEUROVENT®-P and -P-TEMP
These catheter models have an additional blue fixing clip (see image), which is sutured to the scalp together with the white fixing wing.

NEUROVENT®-PTO 2L
[image: image23.png]


 Caution:
The fixing wing may be fixed in place only on the thick part of the catheter hose (8F).

The thin part of the catheter hose (5F) must be laid in a loop Ø approx. 20 mm for strain relief and covered with a plaster.
14. Safety information
· Use of the pressure catheter must take place according to the recognized rules of medical practice (therapy guidelines)!

· CF type (cardiac floating) measuring systems with the [image: image24.png]


 symbol only shall be used.

· The product has not been fitted with protective measures to prevent the patient from being burnt when high-frequency surgical equipment is used.

· The ICP-TEMP-Cable or PTO-Cable must be disconnected from the pressure catheter before using a high-frequency surgical device!

· Because the pressure catheter is not defibrillation protected, the ICP-TEMP-Cable or the PTO-Cable must be separated from the pressure catheter before using a defibrillator!

· High frequency pick-up (e.g. from mobile phones or pagers) may under unfavourable conditions lead to interference in the signal and thus result in incorrect measurements!

· NEUROVENT®-PTO, NEUROVENT®-PTO 2L and NEUROVENT®-TO catheters may only be employed in conjunction with the MPR2 logO or EASY logO monitor!

· Catheter models NEUROVENT®-PTO and NEUROVENT®-TO may be used only with BOLT KIT PTO!
· Catheter version NEUROVENT®-PTO 2L may be used only with BOLT KIT PTO 2L!
· In the event of intentional disconnection (e.g. for care measures, patient transport), the ICP-TEMP-Cable or PTO-Cable must always be disconnected from the pressure catheter first!

· If ICP-TEMP-Cable or PTO-Cable are not connected, the catheter plug must always be closed with the included protective cap to protect it against moisture (see image)!
15.  MRI safety RAUMEDIC® Precision pressure catheters at 1.5 T / 3.0 T

[image: image25.png]


 Note:
Within the scope of non-clinical laboratory tests it was demonstrated that with observance of the following notes and parameters, RAUMEDIC® precision pressure catheters are safe in the sense of MR conditionally in 1.5 T and 3.0 T MRI procedures according to ASTM F 2503-13.
· Different MRI systems with different methods to calculate the scan parameters (e.g. SAR) and individual, technical features of different MRI systems cannot be copied in laboratory tests. The conditions under which the feature MR can be conditionally guaranteed for the RAUMEDIC® precision pressure catheter are listed in point 15.2. MRI parameters. Deviations from these defaults can increase the risk of danger to patients!
· Before the examination, the doctor must assess the risk whether the same information could not also be gained with other imaging methods with a lower risk of danger to the patient.
· In order to safely avoid a displacement of the RAUMEDIC® precision pressure catheter due to the force effect of the spatial magnetic field gradients, the catheter fixation specified for each catheter type in this instruction must be used!

· MRI devices with field strengths other than 1.5 T and 3.0 T may not be used!
· If the region of interest for imaging is near the RAUMEDIC® precision pressure catheter, then the existing image corruption from the object surface (3.0 T spin echo: longitudinal artefact 0.6 mm, transversal artefact 25.6 mm and 3.0 T gradient echo: longitudinal artefact 2.5 mm, transversal artefact 19.6 mm) must be taken into consideration. The scan parameters may have to be optimised for minimal image artefacts or another imaging method may have to be used.

15.1. Use during MRI
· It must be ensured directly before the MRI examination that the RAUMEDIC® precision pressure catheter functions correctly. If this is not the case then the positioned catheter may not be used for an MRI examination!

· The RAUMEDIC® precision pressure catheter required a special positioning geometry for MRI examinations to safely prevent the probe tip from heating up.
· It must be ensured during the MRI examination that, before and during the examination during MRI, there is no disconnection of the RAUMEDIC® accessories, PTO-Cable or ICP‑TEMP‑Cable, used for the respective catheter!
· The low magnetic components of the plugs (on the cable and on the RAUMEDIC® precision pressure catheter) can lead to image artefacts of approx. 70 mm starting from the surface of the object. These components must be place with sufficient distance to the anatomical regions of interest for imaging.
· The cable and the part of the catheter outside the patient must be positioned in coiled condition (loop-diameter approx. 20 cm, complies with approx. 3.0…3.5 coils) at the head of the patient. The cable coil must be set down at the headboard of the patient table in parallel (horizontally) to the static main magnetic field B0 (see image)! This positioning also prevents artefacts in the head of the patient from the electrical and optical plug.
[image: image26.png]


 Caution:
If a deviation from the configuration specified above occurs, there is a clinically unacceptable heating risk under worst case conditions (cable stretched out beside the patient) in MRI 1.5 T and 3.0 T. 

Heating that occurs under required configurations is shown in the Section 17 table “Laboratory data for heating effects during MRI”.

· If the local conditions do not allow the coiled cable to be plugged into the external part of the catheter, the potentially higher heating if a cable is not plugged in (see table Laboratory data for heating effects during MRI) must be considered with appropriate reduction of the scanning time and/or the SAR.
[image: image27.png]


 Note:
With 1.5 T and 3.0 T MRI imaging, plugging the connecting cable in is not essential for safe use if the given instructions are observed. But we recommend plugging in the coiled cable in the sense of a uniform approach for 1.5 T and 3.0 T and the lower heating effects.
15.2. MRI parameters
The MRI examination must be carried out adhering to the following MRI device conditions:

	Static magnetic field B0:
	1.5 T or 3.0 T

	Maximum of the spatial

magnetic field gradient:
	
[image: image28.wmf]B

Ñ

 ≤ 25 T/m (value extrapolated)

	Maximum of the spatial

magnetic field gradient of the product:
	
[image: image29.wmf]B

B

Ñ

×

 ≤ 50 T2/m (value extrapolated) 

	Maximum, theoretically derived, specific absorption rate (SAR) averaged via the head:
	for precision pressure catheters with cable plugged in as per section 15.1.:

< 3.2 W/kg at 1.5 Tesla and 3.0 Tesla (normal mode of operation and initial-stage mode of operation)

	
	for precision pressure catheters without a plugged-in cable:

3.2 W/kg at 1.5 Tesla and 3.0 Tesla (normal mode of operation and initial-stage mode of operation)

	Tested configurations (worst case):
	NEUROVENT®-PTO with Bolt and Cable PTO,

implantation depth ( 3 cm,

Cable PTO ( wound 20 cm, main level of the cable coil parallel to B0, see section 15.1., 

laser cross positioned on the head (head in the isocentre of the body coil)

	
	NEUROVENT®-PTO with Bolt and no Cable PTO,

implantation depth ( 3 cm,

laser cross positioned on the head (head in the isocentre of the body coil)

	
	NEUROVENT®-TEMP-IFD-S with skin tunnel and ICP‑TEMP-Cable,

implantation depth 8 cm (ventricle),

ICP-TEMP cable ( wound 20 cm, main level of the cable coil parallel to B0, see section 15.1., 

laser cross positioned on the head (head in the isocentre of the body coil)

	
	NEUROVENT®-TEMP-IFD-S with skin tunnel and no ICP-TEMP-Cable,

implantation depth 8 cm (ventricle),

laser cross positioned on the head (head in the isocentre of the body coil)

	
	NEUROVENT®-TEMP-IFD-R with Bolt and ICP‑TEMP‑Cable,

implantation depth 8 cm (ventricle),

ICP-TEMP cable ( wound 20 cm, main level of the cable coil parallel to B0, see section 15.1., 

laser cross positioned on the head (head in the isocentre of the body coil)

	
	NEUROVENT®-TEMP-IFD-R with Bolt and no ICP‑TEMP-Cable,

implantation depth 8 cm (ventricle),

laser cross positioned on the head (head in the isocentre of the body coil)

	Max. duration of continuous MRI scan:
	15 min.


16. Indication and warning symbols on the label
	[image: image30.png]



	Sterilized with ethylene oxide
	[image: image31.png]



	Store in a dry place!

	[image: image32.png]



	Not for reuse!
	[image: image33.jpg]



	Warning!

Observe accompanying instruction manual!

	[image: image34.png]



	Only use if packaging is not damaged!
	[image: image35.jpg]



	Date of manufacture

	[image: image36.png]Mo
\\// .




	Store away from sunshine!
	[image: image37.jpg]



	Use by

	[image: image38.png]


0123
	Product meets the basic requirements of EC directive 93/42/EEC. 
	[image: image39.emf]
	The RAUMEDIC® precision pressure catheter is MR conditional

	[image: image40.png]SN





	Serial number
	
[image: image41.png]



	Article number

	[image: image42.png]






	Non pyrogenic
	[image: image43.png]



	Temperature limitation

	[image: image44.jpg]



	Caution!

Observe the operating instructions for important safety-related information
Please note!
	
	


17. Technical data
	Pressure measurement information

	Pressure measurement range:
	-40 to +400 mm Hg (53 kPa)

	Maximum pressure:
	1500 mm Hg (200 kPa)

	Sensitivity:
	5 µV/V/mm Hg ± 1% (37.5 µV/V/kPa)

	Zero point stability in physiolog. saline solution:
	maximum drift of 3.0 mm Hg in 100 hours


	Temperature measurement information (YSI400)

	Measured exit area (initial value range):
	25°C to 45°C

	Measured expanded exit area (initial value range):
	15°C to 25°C

	Accuracy in the range from 25°C to 45°C:
	+ 0.1°C

	Accuracy in the range from 15°C to 25°C:
	+ 0.2°C

	Minimum measurement time:
	10 s

	Temperature measuring method:
	Direct mode, compatible with monitors that support the YSI 400 characteristic curve

	Measuring point/reference body site:
	Measurement of brain temperature


	Oxygen partial pressure measurement ptiO2 information

	Measurement range:
	0-200 mmHg oxygen partial pressure (with MPR2 logO)

	Measurement accuracy:
	1) 0…120 mmHg (
( 3 % of the measured value or 
( 2.50 mmHg (the larger value applies)

2) 120…200 mmHg (
max. 10% of the measured value

	Time drift: (*)
	max. 1.5 mmHg over 5 days (at a ptiO2 of 
[image: image45.wmf]£

10 mmHg)

	ptiO2-sensitive area: (*)
	> 13 mm2 for NEUROVENT®-TO

> 22 mm2 for NEUROVENT®-PTO and -PTO 2L

	
	Temperature range
	Humidity range
	Atmospheric pressure range

	Operating conditions:
	+10°C to +40°C
	30-75% RH
	700-1050 hPa

	Storage conditions:
	+5°C to +40°C
	30-80% RH
	700-1050 hPa

	Transport conditions:
	-20°C to +65°C
	30-85% RH
	700-1050 hPa


	Laboratory data for heating effects during MRI

	Configuration
	Max. temperature increase on the catheter
	Max. temperature increase in the environment
	Temperature difference
	normalised to (WBA)-SAR

	1.5 T *

	NEUROVENT®-PTO with Bolt,

implantation depth 3 cm (parenchymal),

Cable PTO plugged in as per section 15.1.
	2.0°C
	0.4°C
	1.6 K
	2.4 W/kg

	NEUROVENT®-TEMP-IFD-S, skin tunnel,

implantation depth 8 cm (ventricular),

ICP-TEMP-Cable plugged in as per section 15.1.
	0.9°C
	0.6°C
	0.3 K
	3.2 W/kg

	NEUROVENT®-TEMP-IFD-S, skin tunnel,

implantation depth 8 cm (ventricular),

no cable plugged in
	2.1°C
	0.4°C
	1.7 K
	3.2 W/kg

	3.0 T **

	NEUROVENT®-PTO with Bolt,

implantation depth 3 cm (parenchymal),

Cable PTO plugged in as per section 15.1.
	0.6°C
	0.5°C
	0.1 K
	3.2 W/kg

	NEUROVENT®-TEMP-IFD-R, with Bolt,

implantation depth 8 cm (ventricular),

ICP-TEMP-Cable plugged in as per section 15.1.
	0.9°C
	0.5°C
	0.4 K
	3.2 W/kg

	NEUROVENT®-TEMP-IFD-R, with Bolt,

implantation depth 8 cm (ventricular),

no cable plugged in
	4.6°C
	0.2°C
	4.4 K
	3.2 W/kg


*)
Non-clinical examination on the head phantom with body coil with 64 MHz (1.5 Tesla equivalent) RF laboratory system "Medical Implant Test system MITS 1.5", Zurich Medtech AG (software: MITS-DUALBAND 1.2.5.2), 15 minutes of continuous MR scanning. The NEUROVENT®-PTO catheter is elongated for length x und guided in parallel to the static magnetic field B0. Connecting Cable PTO was, as described in 15.1. Use during MRI, coiled and plugged to the catheter. Analogous measurements are performed with the NEUROVENT®-TEMP-IFD-S, skin tunnel, implantation depth 8 cm. The following heating was detected.

**)
Non-clinical examination on the head phantom with body coil, MR Scanner 3 Tesla Magnetom Trio, Siemens (software: Numaris/4, syngo MR A30), 15 minutes of continuous MR scanning. The NEUROVENT®-PTO catheter is elongated for length x und guided in parallel to the static magnetic field B0. Connecting Cable PTO was, as described in 15.1. Use during MRI, coiled and plugged to the catheter. Analogous measurements are performed with the NEUROVENT®-TEMP-IFD-R, with Bolt, implantation depth 8 cm. The following heating was detected.

With regard to the heating effects, the catheters and types of application used for the laboratory tests constitute the worst-case configurations of the NEUROVENT® catheter family shown in the Section 5 table.
18. Accessories
	Cables

	ICP-TEMP-Cable
	094328-001

	Cable LWL
	095657-001

	Cable PTO
	095624-001

	ICP-TEMP-Adapter 
	094323-001

	ICP-TEMP-Adapter 2m
	095226-001

	ICP-TEMP-Adapter Philips/HP
	094047-001

	ICP-TEMP-Adapter Philips/HP 2m
	095216-001

	Monitors

	MPR 1 DATALOGGER
	094474-001/-002

	MPR2 logO DATALOGGER
	095254-001/-002

	EASY logO
	095264-001/-002

	NPS 3
	091656-001

	Application aids

	BOLT KIT CH5
	091868-002

	BOLT KIT CH9
	091688-002

	BOLT KIT PTO
	096026-001

	BOLT KIT - PTO 2L
	096076-001

	DRILL KIT CH5
	091878-002

	DRILL KIT CH9
	091668-002

	BOLT-DRILL KIT CH5
	091888-001

	BOLT-DRILL KIT CH9
	091898-001

	BOLT-DRILL KIT PTO
	092380-001

	Spliceable Tunnelling Sleeve CH8
	090506-002

	Spliceable Tunnelling Sleeve CH12
	090717-001

	NPS2

	NPS2 Datex Ohmeda
	090924-001

	NPS2 Nihon Kohden BSM 8800
	091676-001

	NPS2 SpaceLabs
	091715-001

	NPS2 Spiegelberg CPP
	091738-001

	NPS2 CODMAN
	091978-001

	NPS2 Hellige 
	092617-001

	NPS2 Siemens/Dräger Infinity
	092627-001

	NPS2 Philips/HP
	092637-001

	NPS2 GE/MARQUETTE
	093807-001

	NPS2 Schiller ARGUS PRO
	094613-001

	NPS2 Braun VISICON
	094624-001

	NPS2 Nihon Kohden BSM 41xx
	094716-001

	NPS2 Fukuda Denshi 8000
	095007-001

	NPS2 SpaceLabs 9X496 Module
	095027-002

	NPS2 Criticare nCompass
	095037-001

	NPS2 Petas
	095206-001

	NPS2 Datascope
	095918-001

	NPS2 Fukuda Denshi
	096003-001

	NPS2 Spiegelberg Compliance
	096013-001

	NPS2 Saadat
	096023-001


19. Liability and guarantee information
RAUMEDIC® precision pressure catheters (referred to in the following as catheters) are handled with extreme care during development, component selection, manufacturing and final testing prior to delivery.

The responsibility for the suitability of the type of catheter used and for the correct employment shall be borne by the user. Incorrect handling may lead to functional impairment which in turn could lead to damage and injury. RAUMEDIC® will not be responsible for any injuries of any type nor for any losses or damages, whether these be direct, indirect or consequential losses resulting from use of the catheter outside its intended use as described in the instruction manual.

As a matter of course however, medical complications cannot be ruled out, nor can functional faults and failures of the connected pressure measurement amplification and monitoring systems and the resulting functional faults and failures of the catheters.

Therefore, no guarantee can be provided, nor can it be warranted or guaranteed that no functional faults or failures of the catheters will occur, that the body will not display an immune reaction or that non-medical complications will not occur as a result of the use of the catheter.

Any guarantee or liability for defects which are not immediately detected during initial use of the catheters is excluded.
20. Disposal
After use, the RAUMEDIC® precision pressure catheters must be correctly disposed of in accordance with the applicable instructions for disposal of contaminated hospital waste.


[image: image46.png]


0123

Bibliography:

1) Abdulla W. Intrakranielle Druckmessung. In: Abdulla W. Interdisziplinäre Intensivmedizin. 3. Auflage. München [u.a.]: Elsevier, Urban & Fischer 2007; 658-660 
2





4





1





3





Joining the optical fibre plug connection





measurement head





stylet








Manufacturer: RAUMEDIC AG, Hermann-Staudinger-Str. 2, 95233 Helmbrechts, Germany

zwo-013EN / Rev. Status: 28
Page 2 of 19

_1420957364.unknown

_1502897864

_1420041120.unknown

_1215935050.unknown

