OLYMPUS AU 400/640

Specific Test Parameters

General

	Test Name
	CA
	Type
	Serum
	Operation
	Yes

	
	
	
	
	
	

	
	
	
	

	Sample 
	Volume
	2.0
	µl
	Dilution
	0
	µl

	Reagents
	R1 Volume
	160
	µl
	Dilution
	0
	µl

	
	R2 Volume
	40
	µl
	Dilution
	0
	µl

	
	
	
	

	Wavelength
	Pri
	660
	
	Sec
	800
	

	Method
	
	END
	
	

	Reaction Slope
	
	-
	
	

	Measuring point 1
	First
	0
	
	Last
	15
	

	Measuring point 2
	First
	0
	
	Last
	10
	

	Linearity
	
	%
	

	No-Lag-Time
	
	
	

	Pre-dilution Rate
	
	
	

	
	
	
	

	Min OD
	
	Max OD
	

	
	
	
	
	

	Reagent OD Limit
	
	
	

	
	First L
	-0.100
	
	First H
	2.50
	

	
	Last L
	-0.100
	
	Last H
	2.50
	

	Dynamic Range
	
	
	

	L
	0.0
	H
	25.0
	

	Correlation Factor
	
	
	

	A
	1
	B
	0
	

	On-board stability period
	
	30 Days
	

	
	
	
	

	Value/Flag
	#
	Level L
	#
	Level H
	#
	

	Normal Ranges
	Age L
	
	Age H

	
	Sex
	
	Year
	
	Month
	
	Year
	
	Month
	

	1
	#
	
	#
	
	#
	
	#
	
	#
	

	2
	#
	
	#
	
	#
	
	#
	
	#
	

	3
	#
	
	#
	
	#
	
	#
	
	#
	

	4
	#
	
	#
	
	#
	
	#
	
	#
	

	5
	#
	
	#
	
	#
	
	#
	
	#
	

	6
	#
	
	#
	
	#
	
	#
	
	#
	

	7
	#
	
	#
	
	#
	
	#
	
	#
	

	
	
	
	
	
	
	

	
	
	L
	
	H
	
	

	Panic Value
	
	#
	
	#
	Unit
	mg/dl
	

	
	
	
	
	
	
	

	Calibration Type
	AB
	Formula
	Y=AX+B
	Counts
	#
	

	
	
	
	
	

	
	
	
	Factor
	
	Factor
	

	
	Cal.No
	
	OD
	
	CONC
	
	OD-L
	
	OD-H
	

	Point 1
	*
	
	
	
	*
	
	
	
	
	

	Point 2
	*
	
	
	
	*
	
	
	
	
	

	Point 3
	*
	
	
	
	*
	
	
	
	
	

	Point 4
	*
	
	
	
	*
	
	
	
	
	

	Point 5
	*
	
	
	
	*
	
	
	
	
	

	Point 6
	*
	
	
	
	*
	
	
	
	
	

	Point 7
	*
	
	
	
	*
	
	
	
	
	

	
	
	
	
	

	1-Point Cal. Point
	
	
	
	

	
	
	
	
	

	MB Type Factor
	
	Calibration Stability Period
	
	

	
	
	
	
	


#)
Data entry by the user

*)
Enter calibration or standard value and position
CALCIUM P FS

Order information

Cat. No. 1 1181 .. .. …

Notes

1.
Please refer to the package insert for Calcium P FS for detailed information about the test on the following:
Clinical Relevance

Method and Principle

Composition and Stability of the Reagents

Specimens

Calibrators and Controls

Performance Characteristics regarding



- Measuring Range



- Specificity/Interferences



- Sensitivity/Limit of Detection



- Precision (Reproducibility, Repeatability)



- Method Comparison

Reference Ranges

Literature

2.
The stability of the reagent on board the analyser is 7 weeks provided that contamination and evaporation are avoided.

3. Manufactured by 

DiaSys Diagnostic Systems GmbH 

Alte Strasse 9, 65558 Holzheim, Germany

**This application proposal is for guidelines only. To avoid misinterpretation measured results have to be validated and assessed with caution.


845 1181 10 02 21








June 2010/4

