OLYMPUS AU 400/640

Specific Test Parameters

General

	Test Name
	CREA
	Type
	Serum
	Operation
	Yes

	
	
	
	
	
	

	
	
	
	

	Sample 
	Volume
	10.0
	µl
	Dilution
	0
	µl

	Reagents
	R1 Volume
	140
	µl
	Dilution
	0
	µl

	
	R2 Volume
	35
	µl
	Dilution
	0
	µl

	
	
	
	

	Wavelength
	Pri
	520
	
	Sec
	800
	

	Method
	
	FIXED
	
	

	Reaction Slope
	
	+
	
	

	Measuring point 1
	First
	13
	
	Last
	27
	

	Measuring point 2
	First
	
	
	Last
	
	

	Linearity
	
	%
	

	No-Lag-Time
	
	
	

	Pre-dilution Rate
	
	
	

	
	
	
	

	Min OD
	
	Max OD
	

	
	-0.200
	
	2.50
	

	Reagent OD Limit
	
	
	

	
	First L
	-0.200
	
	First H
	2.00
	

	
	Last L
	-0.200
	
	Last H
	2.00
	

	Dynamic Range
	
	
	

	L
	0.0
	H
	15.0
	

	Correlation Factor
	
	
	

	A
	1
	B
	0
	

	On-board stability period
	
	30 Days
	

	
	
	
	

	Value/Flag
	#
	Level L
	#
	Level H
	#
	

	Normal Ranges
	Age L
	
	Age H

	
	Sex
	
	Year
	
	Month
	
	Year
	
	Month
	

	1
	#
	
	#
	
	#
	
	#
	
	#
	

	2
	#
	
	#
	
	#
	
	#
	
	#
	

	3
	#
	
	#
	
	#
	
	#
	
	#
	

	4
	#
	
	#
	
	#
	
	#
	
	#
	

	5
	#
	
	#
	
	#
	
	#
	
	#
	

	6
	#
	
	#
	
	#
	
	#
	
	#
	

	7
	#
	
	#
	
	#
	
	#
	
	#
	

	
	
	
	
	
	
	

	
	
	L
	
	H
	
	

	Panic Value
	
	#
	
	#
	Unit
	mg/dl
	

	
	
	
	
	
	
	

	Calibration Type
	AB
	Formula
	Y=AX+B
	Counts
	#
	

	
	
	
	
	

	
	
	
	Factor
	
	Factor
	

	
	Cal.No
	
	OD
	
	CONC
	
	OD-L
	
	OD-H
	

	Point 1
	*
	
	
	
	*
	
	
	
	
	

	Point 2
	*
	
	
	
	*
	
	
	
	
	

	Point 3
	*
	
	
	
	*
	
	
	
	
	

	Point 4
	*
	
	
	
	*
	
	
	
	
	

	Point 5
	*
	
	
	
	*
	
	
	
	
	

	Point 6
	*
	
	
	
	*
	
	
	
	
	

	Point 7
	*
	
	
	
	*
	
	
	
	
	

	
	
	
	
	

	1-Point Cal. Point
	
	
	
	

	
	
	
	
	

	MB Type Factor
	
	Calibration Stability Period
	
	

	
	
	
	
	


#)
Data entry by the user

*)
Enter calibration or standard value and position

CREATININE FS

Order information

Cat. No. 1 1711 .. .. …
Notes

1.
Please refer to the package insert for Creatinine FS for detailed information about the test on the following:
Clinical Relevance

Method and Principle

Composition and Stability of the Reagents

Specimens

Calibrators and Controls

Performance Characteristics regarding



- Measuring Range



- Specificity/Interferences



- Sensitivity/Limit of Detection



- Precision (Reproducibility, Repeatability)



- Method Comparison

Reference Ranges

Literature

2.
The stability of the reagent on board the analyser is at least one month provided that contamination and evaporation are avoided.

3. Manufactured by 

DiaSys Diagnostic Systems GmbH 

Alte Strasse 9, 65558 Holzheim, Germany
**Applications developed before December 2003 had been used at a variety of customer laboratories and are validated by use. Software upgrades at specific analysers may require further adjustments as they are out of our control

OLYMPUS AU 400/640

Specific Test Parameters

General

	Test Name
	CREA
	Type
	Serum
	Operation
	Yes

	
	
	
	
	
	

	
	
	
	

	Sample 
	Volume
	10.0
	µl
	Dilution
	0
	µl

	Reagents
	R1 Volume
	140
	µl
	Dilution
	0
	µl

	
	R2 Volume
	35
	µl
	Dilution
	0
	µl

	
	
	
	

	Wavelength
	Pri
	520
	
	Sec
	800
	

	Method
	
	FIXED
	
	

	Reaction Slope
	
	+
	
	

	Measuring point 1
	First
	13
	
	Last
	27
	

	Measuring point 2
	First
	
	
	Last
	
	

	Linearity
	
	%
	

	No-Lag-Time
	
	
	

	Pre-dilution Rate
	
	
	

	
	
	
	

	Min OD
	
	Max OD
	

	
	-0.200
	
	2.50
	

	Reagent OD Limit
	
	
	

	
	First L
	-0.200
	
	First H
	2.00
	

	
	Last L
	-0.200
	
	Last H
	2.00
	

	Dynamic Range
	
	
	

	L
	0.0
	H
	15.0
	

	Correlation Factor
	
	
	

	A
	1
	B
	-0.3
	

	On-board stability period
	
	30 Days
	

	
	
	
	

	Value/Flag
	#
	Level L
	#
	Level H
	#
	

	Normal Ranges
	Age L
	
	Age H

	
	Sex
	
	Year
	
	Month
	
	Year
	
	Month
	

	1
	#
	
	#
	
	#
	
	#
	
	#
	

	2
	#
	
	#
	
	#
	
	#
	
	#
	

	3
	#
	
	#
	
	#
	
	#
	
	#
	

	4
	#
	
	#
	
	#
	
	#
	
	#
	

	5
	#
	
	#
	
	#
	
	#
	
	#
	

	6
	#
	
	#
	
	#
	
	#
	
	#
	

	7
	#
	
	#
	
	#
	
	#
	
	#
	

	
	
	
	
	
	
	

	
	
	L
	
	H
	
	

	Panic Value
	
	#
	
	#
	Unit
	mg/dl
	

	
	
	
	
	
	
	

	Calibration Type
	AB
	Formula
	Y=AX+B
	Counts
	#
	

	
	
	
	
	

	
	
	
	Factor
	
	Factor
	

	
	Cal.No
	
	OD
	
	CONC
	
	OD-L
	
	OD-H
	

	Point 1
	*
	
	
	
	*
	
	
	
	
	

	Point 2
	*
	
	
	
	*
	
	
	
	
	

	Point 3
	*
	
	
	
	*
	
	
	
	
	

	Point 4
	*
	
	
	
	*
	
	
	
	
	

	Point 5
	*
	
	
	
	*
	
	
	
	
	

	Point 6
	*
	
	
	
	*
	
	
	
	
	

	Point 7
	*
	
	
	
	*
	
	
	
	
	

	
	
	
	
	

	1-Point Cal. Point
	
	
	
	

	
	
	
	
	

	MB Type Factor
	
	Calibration Stability Period
	
	

	
	
	
	
	


#)
Data entry by the user

*)
Enter calibration or standard value and position

CREATININE FS

Compensated Method

To use the Creatinine FS Compensated Method on  the Olympus AU 400 and AU 640 the Correlation  factor ’B’ must be changed as follows;

When the units are mg/dL
	Correlation Factor 
	B =
	- 0.3
	


When the units are µmol/L
	Correlation Factor 
	B =
	- 27
	


Note: Ensure that when using the Creatinine Compensated Method  the relevant assigned Calibrator value is selected







845 1711 10 02 21

July 2001/3


