OLYMPUS AU 400/640

General

	Test Name
	IRON
	Type
	Serum
	Operation
	Yes

	
	
	
	
	
	

	
	
	
	

	Sample 
	Volume
	12.0
	µl
	Dilution
	0
	µl

	Reagents
	R1 Volume
	140
	µl
	Dilution
	0
	µl

	
	R2 Volume
	40
	µl
	Dilution
	0
	µl

	
	
	
	

	Wavelength
	Pri
	600
	
	Sec
	700
	

	Method
	
	END
	
	

	Reaction Slope
	
	+
	
	

	Measuring point 1
	First
	0
	
	Last
	27
	

	Measuring point 2
	First
	0
	
	Last
	10
	

	Linearity
	
	%
	

	No-Lag-Time
	
	
	

	Pre-dilution Rate
	
	
	

	
	
	
	

	Min OD
	
	Max OD
	

	
	
	
	
	

	Reagent OD Limit
	
	
	

	
	First L
	-0.1000
	
	First H
	1.500
	

	
	Last L
	-0.1000
	
	Last H
	1.500
	

	Dynamic Range
	
	
	

	L
	5.0
	H
	800.0
	

	Correlation Factor
	
	
	

	A
	1.000
	B
	0.000
	

	On-board stability period
	
	40 Days
	

	
	
	
	

	
	
	
	

	
	
	
	

	Value/Flag
	#
	Level L
	#
	Level H
	#
	

	Normal Ranges
	Age L
	
	Age H

	
	Sex
	
	Year
	
	Month
	
	Year
	
	Month
	

	1
	#
	
	#
	
	#
	
	#
	
	#
	

	2
	#
	
	#
	
	#
	
	#
	
	#
	

	3
	#
	
	#
	
	#
	
	#
	
	#
	

	4
	#
	
	#
	
	#
	
	#
	
	#
	

	5
	#
	
	#
	
	#
	
	#
	
	#
	

	6
	#
	
	#
	
	#
	
	#
	
	#
	

	7
	#
	
	#
	
	#
	
	#
	
	#
	

	
	
	
	
	
	
	

	
	
	L
	
	H
	
	

	Panic Value
	
	#
	
	#
	Unit
	µg/dl
	

	
	
	
	
	
	
	

	
	
	
	
	
	
	

	
	
	
	
	
	
	

	Calibration Type
	AB
	Formula
	Y=AX+B
	Counts
	#
	

	
	
	
	
	

	
	
	
	Factor
	
	Factor
	

	
	Cal.No
	
	OD
	
	CONC
	
	OD-L
	
	OD-H
	

	Point 1
	*
	
	
	
	*
	
	
	
	
	

	Point 2
	
	
	
	
	
	
	
	
	
	

	Point 3
	
	
	
	
	
	
	
	
	
	

	Point 4
	
	
	
	
	
	
	
	
	
	

	Point 5
	
	
	
	
	
	
	
	
	
	

	Point 6
	
	
	
	
	
	
	
	
	
	

	Point 7
	
	
	
	
	
	
	
	
	
	

	
	
	
	
	

	1-Point Cal. Point
	
	
	
	

	
	
	
	
	

	MB Type Factor
	
	Calibration Stability Period
	
	

	
	
	
	
	


#)
Data entry by the user

*)
Enter calibration or standard value and position
IRON FS Ferene

Order information

Cat. No. 1 1911 .. .. …
Notes

1.
Please refer to the package insert for Iron FS Ferene for detailed information about the test on the following:
Clinical Relevance

Method and Principle

Composition and Stability of the Reagents

Specimens

Calibrators and Controls

Performance Characteristics regarding



- Measuring Range



- Specificity/Interferences



- Sensitivity/Limit of Detection



- Precision (Reproducibility, Repeatability)



- Method Comparison

Reference Ranges

Literature

2.
The stability of the reagent on board the analyser is at least 40 days provided that contamination and evaporation are avoided.

3. Manufactured by 

DiaSys Diagnostic Systems GmbH 

Alte Strasse 9, 65558 Holzheim, Germany

**This application proposal is for guidelines only. To avoid misinterpretation measured results have to be validated and assessed with caution.
August 2017/1

