
About The User Manual

Thank you for purchasing our product！
In order to enable you to skillfully operate the device as soon as possible, a detailed user manual is attached. Please make sure to read all the content when installing and using the device for the first time.
To improve the performance and reliability of its parts, the device (including hardware and software) may be changed from time to time, during which, we will try to modify or add contents. Please forgive us as there may still be inconsistency with some descriptions.
We look forward to your corrections in case of any errors and omissions in this manual.

All rights reserved © Shenzhen Biocare Bio-Medical Equipment Co., Ltd.

[image: image1.png]slocare



Information contained in this document is proprietary to Shenzhen Biocare Bio-Medical Equipment Co., Ltd. No part of this publication may be reproduced, stored in a retrieval system, or transmitted in any form, in whole or in part, by any means electronic, mechanical or other, including photocopying and recording, for any purpose without written permission of Biocare.

[image: image132.jpg]


  are trademarks or registered trademarks owned by our company.
	[image: image123.jpg]



	The CE mark is a protected conformity mark of European 

Community. The products herewith comply with the requirements of the Medical Device Directive 93/42/EEC


	[image: image2.png]



	Shanghai International Holding Corp. GmbH(Europe)

Eiffestraße 80

20537 Hamburg

GERMANY


	[image: image3.jpg]



	Shenzhen Biocare Bio-Medical Equipment Co., Ltd.

#16-1, Jinhui Road, Jinsha Community, Kengzi Sub-District, Pingshan New District, 518122 Shenzhen, PEOPLE'S REPUBLIC OF CHINA
Tel: 86 -755 -36615333  Fax: 86-755-27960643

Website: http://www.biocare.com.cn


	CAUTION: In US, Federal law restricts this device to sale by or on the order of a physician.Please read the user manual carefully prior to use.


Explanation of key words
· WARNING
Indicates a potentially hazardous situation which, if not avoided, could result in death or serious injury.

· CAUTION

Indicates a potentially hazardous situation which, if not avoided, may result in minor personal injury or equipment failure.

· EXPLANATION
Indicates other important information besides warning or caution.

Explanation of part symbols

	Symbol
	Explanation
	Symbol
	Explanation

	[image: image124.jpg]


   
	Alternating Current Working
	[image: image4.emf]
	SD Card Slot

	[image: image5.emf]
	Directing Current Working
	[image: image6.emf]
	Patient Cable Socket

	[image: image7.emf]
	Battery Charging
	
[image: image8.png]



	DC Socket

	[image: image9.emf]
	Equipotential Terminal
	[image: image10.emf]
	Please refer to user manual.

	[image: image11.emf]
	USB Port
	[image: image12.emf]
	LAN Port

	[image: image13.emf]
	Type CF Applied Part
	[image: image14.emf]
	Type CF Applied Part including Defibrillation-Proof


Conventions

	Format
	Explanation

	”***”
	Used to quote the texts in the software interface.

	[****]
	Used to quote the shortcut buttons or keys in the software interface.

	TEXT
	Used to quote the referenced chapters or sections in this manual.


Content
IAbout The User Manual


IContent


1Foreword


9Chapter 1 Introduction


101.1 Equipment Overview


111.2 Equipment Description


111.2.1 Front View


111.2.2 Back View


121.2.3 Top View


131.2.4 Left View


131.2.5 Right View


141.2.6 Keyboard View


161.3 Waveform Display


161.3.1 Same screen display


181.3.2 Split-screen display


19Chapter 2 Preparing Recording


192.1 Location Selection


192.2 Installing Recording Paper


192.2.1 Installing rolling paper


202.2.2 Installing Z-fold Paper


222.3 Power Connection


232.4 Patient Cable Connection


242.5 Insert or Remove the SD Card


242.5.1 Insert the SD Card


242.5.2 Remove the SD Card


242.6 Power On/Off


252.7 Network Connection


252.7.1 Connecting to a Wired Network


262.7.2 Connecting to a Wireless Network (Optional)


272.8 Apply Electrodes


292.8.1 Electrodes Attachment


302.8.2 Lead signals formation scheme


31Chapter 3 Entering Patient Information


313.1 Enter Patient Information


323.2 Introduction of Input Method


323.2.1 Standard Character Keyboard


333.2.2 Digital Keyboard


343.2.3 How to input quickly


35Chapter 4 ECG Recording


364.1 Introduction of Sensitivities, Filters, Paper speed


374.2 Recording an ECG


384.2.1 Main operation process to record ECG


394.2.2 Arrhythmia Detection


404.3 Introduction of Recording Mode


414.3.1 Automatic Recording Mode


424.3.2 Manual Recording Mode


434.3.3 Economic Recording Mode


434.3.4 Cycle Recording Mode


434.3.5 Trigger Mode


434.3.6 Upload Mode


444.4 Advanced Mode


444.4.1 Rhythm Mode


454.4.2 Analysis Report Mode


484.4.3 Freeze Mode


49Chapter 5 Setting System Parameters


495.1 ECG Setting


515.2 Print Setting


525.3 Display Setting


525.4 Patient Information Setting


535.5 System Setting


555.6 Remote Report


555.7 Factory Maintain


57Chapter 6 Data Management


586.1 Open an ECG File


586.2 Edit an ECG File


586.3 Delete ECG Files


596.4 Copy and Move ECG files


61Chapter 7 Maintenance


617.1 Main Unit


617.2 Patient Cable


617.3 Cleaning and Disinfection


627.4 Recording Paper


627.5 Battery


647.6 Silicon Rubber Shaft for Printing


647.7 Thermal Printing Head


65Chapter 8 Troubleshooting


658.1 Lead Fault


668.2 Printer Failure


668.3 Keyboard Failure


678.4 Indication of Lead Off


678.5 AC Interference


688.6 EMG Interference


688.7 Baseline Wander


698.8 The ECG machine cannot be turned on


698.9 Paper Feeding Failure


708.10 Battery is quickly charged and discharged


708.11 Wrong Analysis Result


718.12 File Uploading Failure


73Appendix A Package and Accessories


73A.1 Packing List


74A.2 Dimensions and Weight


75Appendix B Technical Specification


75B.1 Specifications


75B.1.1 Main unit


76B.1.2 Recorder Specification


76B.1.3 Other Specification


77B.2 Environment Requirements


79Appendix C Working principle and block diagram


79C.1 Power supply subsystem


80C.2 Acquisition module


81C.3 Control system


83Appendix D List of Interpretation codes and Corresponding Description


89Appendix E Measurement, Diagnosis, Analysis and Assessment of ECG Machine


89E.1 Methods to determine the amplitude of P, QRS, ST and T wave


90E.2 Processing method of isoelectric segment in QRS complex


90E.3 Low incidence heart disease not included in testing and diagnosis database


91E.4 ECG diagnosis categories and the number of ECG test of each category


92E.5 The smallest waveform identified by the device and the stability of measurement when noise exists


92E.6 Low incidence cardiac rhythm not included in the ECG rhythm test database


92E.7 ECG rhythm diagnosis categories and ECG test number of each category


93E.8 Sensitivity regularly test instructions


94E.9 Distortion test


95Appendix F EMC-Guidance and manufacture’s declaration


95F.1 Guidance and manufacturer’s declaration – electromagnetic emission


96F.2 Guidance and manufacture’s declaration-electromagnetic immunity for all EQUIPMENT and SYSTEMS.


97F.3 Guidance and manufacture’s declaration-electromagnetic immunity for EQUIPMENT and SYSTEMS those are not LIFE-SUPPORTING


98F.4 Recommended separation distance between portable and mobile RF communications equipment and the EQUIPMENT or SYSTEM for EQUIPMENT or SYSTEM that are not LIFE-SUPPORTING





Foreword
· CAUTION
· This ECG machine shall be used by qualified health professionals in the medical units, who need to analyze the ECG waveforms and give diagnostic results.
· In sequence to use this ECG machine correctly, safely and effectively, please read through the user manual carefully.
· Safety Information
· WARNING
· Avoid using and storing in the places with sulfur, salt, alkaline gas or with risk of gas leakage.
· Avoid using in the places with anesthetic gases, flammable gases such as oxygen, hydrogen or other flammable chemicals, or it may cause explosion or fire.
· Select a room with intact infrastructure (good power supply system and grounding facilities).
· Be cautious when the patient is connected with more than one instrument, because the total leak current may be harmful to the patient. Devices in compliance with the standard of IEC60601-1 are allowed to be connected to this ECG machine, and the equipotential points of all the connected devices should be connected reliably. (The equipotential point and the protection ground of this ECG machine have been connected). Total leak current should be measured by the users to determine that it meets the requirement and can be used after connection.
· All the analog and digital equipment that is connected to this ECG machine in the patient environment has to be in compliance with the standard IEC60601-1; All the analog and digital equipment that is connected to this ECG machine out of patient environment has to be in compliance with other national safety standards (IEC or ISO safety standards); the composition system should be in compliance with the standards of IEC 60601-1-1.
· If a cardiac pacemaker is implanted in the patient, it might affect the accuracy of analysis result. In this case, the doctor is suggested to identify and analyze according to the waveforms.
· When the equipment is used simultaneously with cardiac defibrillators, avoid contacting with patients or hospital beds. All the electrodes connected and unconnected to patients as well as patients themselves do not have to be grounded. Do not use other electrical stimulators at the same time. If needed, it should be a professional technician to carry out the operation.
· Chest and limb electrodes along with the device in the packing box could not meet the requirements of defibrillation polarization recovery time (however, they are essential accessories of ECG), should not be used immediately for reliable measurements and diagnostics after defibrillation. To ensure proper defibrillator protection, use only the recommended disposable electrodes (Name: Skintact; Type, RT-34), patient cable and electrode adapters by our company. To ensure the protection of defibrillator discharge, use the patient cable with defibrillation-proof by our company.
· When the ECG machine is used together with a defibrillator or other electrical stimulators (like high-frequency surgical devices), we recommend using disposable chest electrodes. Otherwise the patient may get serious injury by using mental electrodes.
· During defibrillation, the device can detect the discharge of defibrillator, and process automatically, and then quickly recover the waveforms.
· Keep the ECG machine electrodes away from the electrodes of high-frequency electrosurgical units. Ensure the resistance between the electrosurgical unit and patient body is as low as possible. If necessary, the disposable electrodes can be used because of its larger contact area on the human body, which can keep the high-frequency current density in an acceptable range.  

· When the relevant packaging material, including depleted batteries and scrapping products are disposed, please follow the local laws; the user should properly follow requirements of local laws, and recycling laws.
· CAUTION
· Avoid contacting with water or other liquids, and avoid using and storing in spaces with too large barometric pressure, humidity and temperature beyond the prescribed standards, poor ventilation, or with excessive dust.
· The ECG machine should be placed on the flat horizontal table and avoid strong vibration and mechanical shock while moving.
· The frequency of AC power supply and system voltage should comply with the requirements. More importantly, the current capacity should be sufficient.
· The instrument should not be surrounded by high-voltage cables, ultrasound equipment, electrotherapy machines and other high-power equipment.
· To more accurately record the ECG, the equipment shall be placed in a quiet and comfortable environment.

· The circuit of applied parts works based on floating ground and meets the safety standards in IEC60601-1 CF Type. It can be used in acquiring the body surface ECG signals, but cannot be applied to the heart directly.
· Turn off the ECG machine if an accident happens.
· Please clean and disinfect reusable electrodes with medical alcohol before usage.
· The conductive parts of electrodes and connectors (including neutral electrodes) on the ECG machine should not be in touch with other conductive parts.
· Do not press the buttons with sharp or hard articles or it may cause permanent damage to the buttons.
· Do not make any modifications to this ECG machine.
· Perform regular maintenance and inspection for this ECG machine and all its accessories (at least once every six months).
· The maintenance and repair of this ECG machine should be performed by experienced technicians. When there is any functional abnormality, it should be clearly identified to prevent the ECG machine from running with fault.
· The electrical schematic diagrams and parts listed are only provided to a qualified repair station or technicians recognized by the company.
· General Operation Precautions

Before operation:

· CAUTION
· Make sure the ECG machine is in a complete and normal condition and whether the recording paper is sufficient.
· Make sure the temperature and humidity of operating environment comply with the requirements.
· Do not operate the ECG machine in an environment with X-ray equipment, ultrasound scanners, or similar equipment. This equipment may interfere with the unit. If necessary, please power off the mentioned equipment or move the ECG machine to an environment without interference.
· Make sure all lead wires and acquisition module are connected correctly and are kept away from the AC power cable.
· Make sure the equipotential cable of this ECG machine is reliably and properly connected.
· Make sure the power cable is properly connected with the ECG machine and is not intertwisted with other cables or wires.
· Put the lead wires in good order before connecting them to the electrodes.
· Make sure the electrodes are in good contact with skin. Please refer to Apply Electrodes for details.
· Make sure the patient’s skin contacted with the electrodes has been well pretreated.
· Clean the stain on the electrodes with alcohol whenever the electrodes are contaminated.
· Make sure clip-electrodes touch fully with skin and tightly enough.
· Please install the ECG machine near an AC power outlet. Cut off the power supply immediately when there is an emergency.
· If the patient is nervous, please explain to the patient that ECG examination is easy and no harm.
· Please keep the patient silent and motionless.

· Use wide hospital beds and keep the patient from touching the metal parts of the hospital bed, which may cause interference in ECG waveform recording.
· Keep the Examination room silent and comfortable.
· WARNING
· All circuits that come in contact with the patient directly should be examined closely.
· When using the battery as the power supply, please check the voltage and condition of the battery first and make sure the battery fully charged. For new batteries, please discharge and charge it fully before usage.
· Use only 3-core power cable when using AC power, otherwise hazard of electric shock to the patient and operator cannot be completely eliminated. If the power cable is not working, only the built-in battery can safely power the ECG machine.
· Make sure equipotential connection is complete and reliable, or else only use the built-in battery.
In operation:
· WARNING
· The physician should observe the patients closely without leaving during the operation. If necessary, turn off the ECG machine and remove the electrodes to ensure patient’s safety.
· Prevent the patients from contacting the other parts of the ECG machine or other conductors except for the electrodes.
After operation:

· CAUTION 
· Please return to the main interface before turning off the ECG machine.
· Remove the electrodes gently and do not pull the lead wires emphatically.
· Clear up the ECG machine and all the accessories for trouble-free operation of next use.
About LCD screen
· CAUTION
· Do not place any heavy objects onto the LCD screen or strike it, otherwise it could cause damage to LCD screen.
· When not using it, please put it away or have a cover on it. Keep it away from water.
· Do not use LCD screen with excessive force.
· If the cursor cannot be correspondent to the touching point, you need to carry out the touch screen calibration.
· No water or soft sticky material is allowed on the touch screen, otherwise it would cause touch error.

About lithium batteries

· WARNING
· Only the authorized installation or service engineer can open the battery cover and replace the battery; do use the same type of rechargeable lithium battery provided by our company.
· The positive and negative terminals of the batteries cannot be reversed, or it could cause an explosion.
· Do not connect the two polarities of the battery with metal wires. Otherwise, there will be the hazard of fire.
· Do not use the battery near a heat source or in an environment with temperature up to 60 ℃; do not heat the battery or throw it into the fire.
· Keep the battery away from water; do not drop the battery into the water.
· Don’t press any metal into the battery; Do not hammer or beat the battery or use other ways to damage the battery, otherwise it will cause heat, smoke, deforming or burning, which is very dangerous.
· When you find battery leakage or its emitting unpleasant odors, please get away from it immediately. If the fluid leaks onto the skin or clothes, wash with clean water at once. If the electrolyte enters the eyes, do not rub the eyes, wash with clean water immediately, and then go to see a doctor.
· The user needs to check the battery working status regularly. When the battery reaches the end of its lifetime, when it smells, deforms, discolor, contorts, the users should stop using and dispose of it according to local regulations.
· EMC Considerations

This ECG machine conforms to the IEC60601-1-2, a safety standard for medical electronic devices or systems. However, the electromagnetic environment exceeding the limit or level defined by the standard IEC60601-1-2 will introduce the unwanted interference to the ECG machine, disable its intended functions or it will compromise its intended performance. Thus, if there is any discrepancy with this ECG machine compared to its intended functions during operation, please do not use it any longer until the adverse effect is identified and eliminated. The appropriate preventing measures are given below by this manual for such cases:
· Influence of radiated electromagnetic wave:
The use of a mobile phone may affect this ECG machine. Instruct all the people around to turn off their mobile phones or mini-radio devices when any medical electronic device is in use.
· Influence of impact and conductive electromagnetic waves:
The high frequency noise produced by other devices can be introduced into this ECG machine through the alternating current socket. Please identify the noise source first, and if possible, stop the working of related devices. If they are not allowed to be stopped, measures such as application of noise abatement device should be taken to minimize the influence.
· Influence of static electricity:
The static electricity in a dry environment (indoor) may affect this ECG machine, especially in winter. Please humidify the indoor air or pre-discharge the static electricity on the cable and the electrocardiogram recording personnel prior to using this ECG machine.
· Influence of thunder and lightning:
A thunder and lightning strike nearby may cause voltage surge in this ECG machine. You can unplug the power supply and run the ECG machine using its internal battery in case of any danger.
Instrument classification
	Methods
	Class

	By Type of Protection Against Electric Shock
	Class I, internal power supply

	By Degree of Protection Against Electric Shock
	Type CF applied part

	By Degree of Liquid Proof
	Ordinary equipment (enclosed device without liquid proof)

	By Level of Protection Against Explosion
	This equipment is unsuitable for use in an environment with air, oxygen or nitrous oxide mixed with flammable anesthetic gas.

	By Mode of Operation
	Continuous operation equipment


Maintenance Warranty
Our company guarantees the new instrument on the material and technological qualification for this product within 18 months and the accessories within 6 months since purchasing day, while consumables are not covered by the guarantee in principle. This guarantee is also inapplicable to the products undergoing any modification, disassembly, refitting or self-repairing without permission of our company, as well as the products damaged by accidents, fire disaster, thunder and lightning, flood and other disaster, intentional damage, improper installation and improper usage. 
· CAUTION
· For all dated reference documents in this manual, its subsequent amendments (excluding corrections) or revisions do not apply to this manual; for undated reference documents, the latest version applies in this manual.

· Due to product improvement, the content of this User Manual may differ from the product you purchased, which will not affect the usage, please operate according to the actual functions of the product.
· This manual is subject to change without prior notice. We apologize for any inconvenience caused.

Chapter 1 Introduction
Composition
The ECG machine is mainly composed of the main unit and accessories. The accessories include patient cables, limb electrodes, chest electrodes, and power adapter etc.

Scope of application

This ECG machine is used to extract the electrocardio complex from the human body to conduct waveforms and rhythm analysis for clinical diagnosis and research.
Intended Use
· The diagnostic applications include: check the cardiac abnormalities of the general population; detect the chest pain in patients with acute myocardial ischemia and myocardial infarction, and evaluate the patients with arrhythmias;

· Suitable for: adults（older than 12 years old）, pediatrics(age between 29 days to 12 years old), and neonates(infants born less than 28 days after 37 weeks to 44 weeks of pregnancy);
· Used in: hospitals, clinics;

· The automatic analysis program of this ECG machine focuses on the high sensitivity of detecting high-risk patients with cardiac abnormalities.
1.1 Equipment Overview
3-channel ECG machine:
[image: image15.jpg]



6-channel ECG machine:
[image: image125.jpg]




	Number
	Name
	Description

	1
	[image: image16.jpg]



	Power the ECG machine on or off.

	2
	Display Screen
	Display the waveforms, patients’ information and the device status.
6-channel ECG machine is equipped with a touch screen and 3-channel ECG machine is not.

	3
	Paper Magazine
	Place recording paper here.

	4
	Paper Magazine Button
	Press to open the paper magazine for loading or unloading the recording paper.

	5
	3-channel：[image: image17.jpg]



6-channel：[image: image18.jpg]



	Start or stop printing the ECG waveforms and report.

	6
	[image: image19.jpg]



	Press to input patient information.


1.2 Equipment Description
1.2.1 Front View
3-channel ECG machine:                                    6-channel ECG machine:
[image: image20.jpg]



1.2.2 Back View
3-channel ECG machine:                                   6-channel ECG machine:
[image: image21.jpg]9

9

—

% =
@*[ )
2 1




	Number
	Name
	Description

	1
	Equipotential Terminal
	Connect to the equipotential cable.

	2
	Power Supply Socket
	Connect to the AC power adaptor.


1.2.3 Top View
3-channel ECG machine:                               6-channel ECG machine:
[image: image22.jpg]00000000000
00000000000
0000000CID0!

10 9 8





	Number
	Name
	Description

	1
	Handle
	Help the user carry the ECG machine easily.

	2
	[image: image23.jpg]



	Power the system on or off.

	3
	[image: image24.jpg]



	Press to input patient information.

	4
	[image: image25.jpg]



	Press to set recording mode and recording format.

	5
	[image: image26.jpg]



	Press Up/Down/Right/Left button to select a menu or an option.

Press “Enter” to confirm, or select an option in a submenu.

	6
	3-channel：[image: image27.jpg]



6-channel：[image: image28.jpg]



	Start or stop printing the ECG waveforms and report.

	7
	[image: image29.jpg]



	Press to switch between main interface and main menu.

	8
	Keyboard
	Enter data into the system.

	9
	Power Indicator
	From left to right: AC power, battery and battery charging.

	10
	Paper Magazine
	Place recording paper here.


1.2.4 Left View
3-channel ECG machine:                             6-channel ECG machine:
[image: image30.jpg]



	Number
	Name
	Description

	1
	Paper Magazine
	Place recording paper here.

	2
	Paper Magazine Button
	Press to open the paper magazine for loading or unloading the recording paper.


1.2.5 Right View
3-channel ECG machine:                             6-channel ECG machine:
[image: image31.jpg]



	Number
	Name
	Description

	1
	Patient Cable Socket
	Connect to the patient cable.

	2
	SD Card Slot
	Insert a SD card.

	3
	USB Port
	Insert a U disk or connect to a bar scanner.

	4
	
	

	5
	LAN Port
	Connect to LAN cable.


1.2.6 Keyboard View
[image: image32.jpg]24
23
22
21

20

|T|ﬂ

P>

1 2 3 4 5 6 7 9 10
n(l F ||<Leaa|.|.Lead>||(| |(| nn}lz ()= @('W
=)OCHCOE)E)OCEE)C)DEL
=) OFDEEOOOEOOCEE
1EEE ICOE)CE)E)EIEE)C ) el
= D000 OELE®EL
) )@= EEE)

19 18

17

16 15

)|
_5/\_|

14

\ .
iy




	Number
	Name
	Description

	1
	Paper Feed
	Press to feed paper. It would stop automatically when it detects the black mark at the bottom of the paper, so that the report can be easier to read.

	2
	Marker
	During printing, press to mark abnormal waveforms by labeling black dot at the bottom of the paper.

	3
	Characters Area
	Press to enter letters, numbers and punctuations

	4
	Lead
	In manual mode, switch among different lead groups.
In split-screen mode, switch among different screens.

	5
	Freeze
	Freeze the waveforms of previous 300 seconds. You can select any waveforms in these 300 seconds to analyze and print.

	6
	Sensitivity
	Switch among different sensitivities.
For 10/5 mm/mV, the sensitivity of limb lead is 10 mm/mV and that of chest lead is 5 mm/mV;

For 20/10 mm/mV, the sensitivity of limb lead is 20 mm/mV and that of chest lead is 10 mm/mV.

	7
	Filter
	Set low-pass filter, baseline wander filter and AC filter.

	8
	Print Speed
	Switch among different printing speeds.

	9
	1mV/Copy
	1.  In manual mode, print the calibration waveforms of 1mV to check the current sensitivity.
2.  In auto mode, copy the previous report.

	10
	Exit
	Exit the current interface and go back to previous one. It is invalid when the ECG machine is already in the main interface.

	11
	Delete
	Delete previous inputted character.

	12
	Enter
	Press “Enter” to confirm, or select an option in a submenu.

	13 and 21
	Shift
	It is not supported in English operation.

	14
	Arrows
	Navigate to select a menu or an option.

	15
	Page Up/Page Down
	Select the previous page or the next page to review more information.

	16
	Screen Key
	Switch between main interface and main menu.

	17
	Upload
	Upload files to the ECG management system on PC.

	18
	Alt
	/

	19
	Fn
	
Mute or unmute the system in the main interface.

	20
	Ctrl
	/

	22
	Caps Lock
	Switch between uppercase and lowercase letters inputting.

	23
	Tab
	Press to move cursor.

	24
	Esc
	Exit the current interface and go back to previous one. It is invalid when the ECG machine is already in the main interface.


1.3 Waveform Display
· EXPLANATION
Screen display may slightly differ from the product you purchased, which will not affect your usage. Please operate according to the actual functions of the product.
In mode of same screen display, 12 lead waveforms will be displayed on one interface.

In mode of split-screen display, 12 lead waveforms will be displayed on several interfaces, which make it possible to show the waveforms details more clearly.

Select [Display], set the display format and lead format.

1.3.1 Same screen display
Waveforms in same screen, 3X4 lead format:
[image: image33.png]Auto, 3%4

10 mm

| n |
i VR - " 1 n A Ava 1) J‘;
|
1 t
J e S J





	Waveforms in same screen,
 6X2 lead format:
	Waveforms in same screen,
 12X1 lead format:

	[image: image34.jpg]avR

avL

avF

Vi

vz

\Z3

V3

V5

V6





	[image: image35.jpg]n
avR
avL
avF

Vi

vz

V3

v4
V5
3







Cabrera lead mode is available in this ECG machine, in which the leads order is aVL, I, -aVR, II, aVF, III, V1 to V6. –aVR is the opposite direction of aVR.
In the menu [ECG Setting], select [Lead Mode], you can switch between “Standard Leads” and “Cabrera”.
Lead waveforms of Cabrera are shown as below:
[image: image36.png]. 065 | 141016003, , , Years ) [ Cabrera @<=

iia A e ns
i 1 1 I ol I
o 4 A A | o I | I
1 A | I





· EXPLANATION
Waveforms cannot be split-screen displayed in Cabrera lead mode.

1.3.2 Split-screen display
Waveforms in split-screen, 6X2 lead format:
[image: image37.jpg]vi

V2

ad

Le

V3
V4
V6





For both 3X4 and 6X2 lead format, please press [image: image38.jpg]


 to change displayed waveforms.
Chapter 2 Preparing Recording
2.1 Location Selection
Please refer to the caution of foreword.

2.2 Installing Recording Paper
· EXPLANATION
Installation of recording paper may slightly differ from the product you purchased, which will not affect your usage, please operate according to the actual functions of the product.
2.2.1 Installing rolling paper
See the descriptions below for how to install the rolling paper into the device:
[image: image39.png]



1. Press the button and open the paper magazine.
2. Pull out the cover of the paper magazine to separate it from the device.
3. Take out the paper shaft from the paper magazine.
4. Insert the paper shaft into the paper.
5. Put the paper into the paper magazine, pay special attention to the direction of paper.
6. Pull the paper out from the outlet of paper magazine and put back the upper cover, and then the recorder is ready.
2.2.2 Installing Z-fold Paper
See the descriptions below for option 1 to install the z-fold recording paper into the device:
[image: image40.jpg]



1. Press the button and open the paper magazine.
2. Pull out the cover of the paper magazine to separate it from the device.
3. Put the paper into the paper magazine, pay special attention to the direction of paper. 
4. Pull the paper out from the outlet of paper magazine and put back the upper cover, press[image: image41.jpg](&)



, and then the recorder is ready.
See the descriptions below for option 2 to install the z-fold recording paper into the device:

[image: image42.jpg]



1. Press the button upside.
2. The cover of paper magazine opens.
3. Push the recording paper fitly into the paper magazine, pay special attention to the position of black mark, and then pull the paper from the outlet of paper magazine.
4. Pull out the paper before putting it back to the magazine and close the magazine with hands, press[image: image43.jpg](&)



, and then the recorder is ready.

· Caution

· Please make sure recording paper is installed fitly and straight, otherwise paper may be stuck.
· If paper is absent, has been used up or not in place, alarm will appear on the main screen and the machine does not print.
2.3 Power Connection

[image: image44.png]



1. Please plug AC power cable and equipotential cable into the equipment.
2. Please connect AC power cable to power adaptor.
3. Please connect 3-core end of power adaptor with power supply socket of the room, and connect equipotential cable with equipotential terminal in the room.
· EXPLANATION
The ECG machine is equipped with a built-in rechargeable battery and any extra installations are not required. Please check the battery capacity before usage.
· caution
· When the ECG machine is operated together with other medical equipment, please use the accompanying equipotential cable and connect the equipotential terminal of the ECG machine altogether with that of the other equipments so as to protect the patient from possible electric shock due to current leakage from those equipments.
· Equipotential cable must be connected between the equipotential terminal of the ECG machine and the room. Do not connect the equipotential cable to a conductive water pipe or the other pipes. Otherwise it may cause electric shock hazard to the patients.
2.4 Patient Cable Connection
[image: image126.jpg]


[image: image127.png]


 
[image: image45.png]



As the picture show, connect the patient cable into the ECG machine. The patient cable is composed of acquisition module and lead wires.
· caution
Do not use any other patient cable except the supplied one. The patient cable socket is exclusively used for connecting the patient cable and do not use it for other purposes.
2.5 Insert or Remove the SD Card
2.5.1 Insert the SD Card
1. Insert a SD card into the SD card slot as shown below.
2. Push to insert a SD card into the SD card slot until a “clicking” sound.
[image: image128.jpg]



[image: image46.png]Card_Slot




2.5.2 Remove the SD Card
1. Push the SD card forward gently until a “clicking” sound.
2. Release the SD card, and then SD card eject about 5 mm.
3. Pinch the SD card away from the slot by nails.
· CAUTION
· If resistance happens when inserting SD card, check whether there is foreign bodies in the slot or the direction of SD card is correct.
· Do not throw or bend the SD card, or it may be broken.
2.6 Power On/Off
Press and hold  [image: image47.png]


 for three seconds to power on/off the ECG machine.

The ECG machine enters standby mode if it is not in use for a certain duration. Set the time duration in [Standby Time] of [System Setting]. Press any key or touch the screen to exit standby mode.

The ECG machine will shut down automatically if it is not in use for a time duration. Set the time duration in [Auto Power Off] of [System Setting].
· EXPLANATION
The ECG machine will not enter standby and auto power off in Auto-trigger mode and cycle mode.
2.7 Network Connection
· EXPLANATION

· The machines include standalone version and network version. Only network version machines are equipped with network features. If network features are needed, please contact our customer service.
· In the ECG machine of network version, network connection may slightly differ from the product you purchased, which will not affect your usage, please operate according to the actual functions of the product.
· In order to realize the data transmission, you need to install ECG management system software and drivers on the computer (Please refer to the user manual of ECG management system). If ECG management system is needed, please contact our customer service.
· CAUTION
In the data transmission, if the ECG machine warns “Network connect failed”, please reset the network.
2.7.1 Connecting to a Wired Network
1. As shown in the following figure, the cable network system is composed of the ECG machine, the switchboard and the server.
2. Set [Transfer Protocol] in [System Setting] according to the protocol of the server. Two options are: TCP and FTP.
3. Set the IP address, subnet mask and gateway in [Wired Network] of the ECG machine. If the IP address is within the same network segment of the server, the subnet mask and gateway shall be as the set value of the server. If the IP address is not within the same network segment of the server, subnet mask and gateway of the [Wired Network] shall be set according to actual situations, but make sure the specified gateway do support the data transmission between the two network segments.
4. Set the IP address and port number in the menu [Server Setting] to the IP address and port number of the server.
5. When the cable network is connected successfully and able to communicate with the server, the system interface will display the icon[image: image48.jpg]


.
[image: image129.bmp][image: image130.png]


[image: image131.emf][image: image49.jpg]ECG machine

vl o
O





2.7.2 Connecting to a Wireless Network (Optional)
1. As is shown in the following figure, the wireless network system is composed of the wireless AP on the ECG machine, the switchboard and the server.
2. Set [Transfer Protocol] in [System Setting] to the protocol by the server. Two options are: TCP and FTP.
3. Set the SSID and password (can be set at will) of the wireless AP and use automatic channel if there is no special requirement.
4. Set the IP address and port number in the menu [Server Setting] as the IP address and port number of the server.
5. Set the SSID, password and Security in [WIFI Network] to the set value and type of the wireless AP. The encryption method should be set up according to that of wireless AP.
6. In [WIFI Network], the setting of the DHCP service can be enabled or disabled according to the requirement. When DHCP service is enabled, it is unnecessary to set IP address, subnet mask and gateway manually. When it is disabled, it is necessary to set IP address, subnet mask and gateway with respect to the setting mode of wired network.
7. After WIFI network is connected, the system interface will display the icon[image: image50.jpg]


.
[image: image51.jpg]ECG machine

=

oeoe
&) TaTAT





2.8 Apply Electrodes
Before attaching electrodes to patient, wipe skin oil off by using medical alcohol to all the positions where electrodes are to be attached, and then apply ECG gel on the skin. After that, place electrodes to the right position. If using the vacuum ball electrodes, you should apply ECG gel on the electrodes, and then pinch the suction ball to make sure the electrodes contact with skin tightly.
· Caution
· Proper electrode attachment is vital for obtaining accurate ECG waveforms; therefore, please ensure good contact between the skin and electrodes.
· Do not use the new electrodes and the used ones at the same time. Replace all electrodes together when any one of them is supposed to be replaced.
· Do not use disposable electrodes more than one time.
· Confirm the disposable electrodes are within the valid period.
· Use the disposable electrodes as soon as possible after opening the package (generally within 7 days).
· Electrodes or conducting point of patient cable shall not be in contacting with any other metal part or conductor.
· Avoid the electrodes to be dragged by the lead wires.
· Make sure the patient’s skin contacted with the electrodes has been well pretreated.
· Clean the stain on the electrodes with medical alcohol whenever the electrodes are contaminated.
· Make sure metal electrodes of limb electrodes touch fully with skin and tightly enough.
· Make sure adjacent electrodes and ECG gel, especially chest ones, are not contacted with each other.
· If the examination involves a short period of time, if ECG gel is unavailable, please wipe the skin with medical alcohol to keep the skin clean and moist, and then rapidly attach the electrodes.
· It is not allowed to use saline water as substitute to ECG gel. The saline water will cause corrosion on the electrodes.
· For chest and back application of pediatric, disposable electrodes are suggested.
· Electrodes shall be properly stored. When electrodes have been used for a certain period, they may become corroded and oxidized at the surface. Whenever this happens, the electrodes must be replaced.
· Do not mix electrodes of different types and manufacturers. Do not use re-useable electrodes and disposable ones together, or it will affect the recording.
· Please use our company’s or authorized electrodes to make sure qualified ECG.
2.8.1 Electrodes Attachment
	Limb Electrodes Placement

	
	IEC
	AHA 
	Description
	Figure

	
	R Red
	RA White
	Right Arm
	[image: image52.jpg]




	
	L Yellow
	LA Black
	Left Arm
	

	
	N Black
	RL Green
	Right Leg
	

	
	F Green
	LL Red
	Left leg
	

	Standard 12-lead Attachment

	
	IEC
	AHA
	Description
	Figure

	A
	C1 Red
	V1 Red
	Fourth inter-costal space at right sternal border.
	

	B
	C2 Yellow
	V2 Yellow
	Fourth inter-costal space at left sternal border.
	

	C
	C3 Green
	V3 Green
	Equidistant between B and D
	

	D
	C4 Brown
	V4 Blue
	Fifth inter-costal space at left mid-clavicles line
	

	E
	C5 Black
	V5 Orange
	Left anterior auxiliary line at the horizontal level of D
	

	F
	C6 Purple
	V6 Purple
	Left mid- auxiliary line at the horizontal level of D
	


2.8.2 Lead signals formation scheme
[image: image53.jpg][ P ey ——
——=e
() (D)
SE’Q =]
() )





Chapter 3 Entering Patient Information
3.1 Enter Patient Information




You can enter patient name, gender, ID Number etc. through the way mentioned above. Refer to Patient Information Setting to get more detail information.
Input modes of ID coding include:
Automatic Coding：ID code is automatically generated by the system when admitting a new patient and an ID code will be automatically increased each time you press it.
Manual Coding: you can input numbers and letters according to his own demand.
Barcode Scanner: you can scan the bar code directly using the scanner to generate ID code. Press [ID Number] and perform a scan to generate the ID code.
[image: image54.png]



· EXPLANATION
When Barcode Scanner is selected as the input mode, the soft keyboard will not pop. For the usage of the barcode scanner, please refer to the user manual of barcode scanner.
· caution
· Improper patient information may cause misdiagnosis. Please check the information for each new patient.
· Please avoid the same ID number for two patients. Otherwise, it may cause file lost or file mistake.
3.2 Introduction of Input Method
You can input through either software keyboard or physical keyboard. The software keyboard on the interface is as following:
3.2.1 Standard Character Keyboard
[image: image55.jpg]2 e e s 7 =
e -

a |l s d k H I

Caps || z

i En | a2 H PgUp H Pgbn





	Number
	Name
	Description

	1
	Characters Area
	Input letters or punctuations.

	2
	Delete
	Delete previous inputted character.

	3
	Cursor Moving
	Move the position of cursor on the interface.

	4
	End
	Confirm the current input.

	5
	Page Down
	/

	6
	Page Up
	

	7
	Space
	Input null characters.

	8
	Symbols
	Switch to symbols pad to input kinds of symbols.

	9
	Caps
	Switch between uppercase and lowercase letters inputting.


3.2.2 Digital Keyboard
[image: image56.jpg]~
w

t

=

[End |





	Number
	Name
	Description

	1
	Characters Area
	Input numbers or punctuations.

	2
	Cursor Moving
	Move the position of cursor on the interface.

	3
	Delete
	Delete previous inputted character.

	4
	End
	Confirm the current input.

	5
	Space
	Input null characters.


3.2.3 How to input quickly
	Step
	Operation
	Button

	1
	Select the item to be operated.
	[image: image57.jpg]




	2
	Confirm and begin inputting. Software keyboard will appear but physical keyboard is recommended on non-touch screen.
	[image: image58.jpg]




	3
	Switch to next item to be operated
	[image: image59.jpg]




	4
	Exit from inputting.
	[image: image60.jpg]




	/
	Switch between uppercase and lowercase letters inputting.
	[image: image61.png]




	/
	Delete characters one by one.
	[image: image62.jpg]





· EXPLANATION
Press [image: image63.jpg]


 to enter input status, the current character strings (if any) will be selected, and press [image: image64.jpg]


 to delete characters one by one from the end of the character string. If you directly input character string, the original selected character string will be cleared.

Chapter 4 ECG Recording
After the ECG machine has been powered on and all the leads are well connected, the following main interface will be displayed. The recording is ready.

[image: image65.png]1 2 3 4 5 6 7 8

Y @1660) = _sn male 28 years) [ 3xa | m &= & ¢ @as
d250Hz Ho.01 ‘Omm/mV  50mmis. Alarm 15:20
[ [|aw ]"‘ v
o= IES ]Vz fve
[ [|ave Ml At nve
= o - ] anll 5] & =
Menu Rhythm Report Freeze (Gain Filter Display Order Print

10 9




	Number
	Name
	Description

	1
	Lead Status
	Display the location and state of the electrode in the human body. Selected for a larger view. If the lead wires are not connected properly, such as electrode falls off, the corresponding electrodes on the view will flash to alarm.

	2
	Heart Rate Icon
	Display patient’s heart rate.

	3
	ECG Setting
	Display current filter, sensitivity and paper speed.

	4
	Patient Information
	Display patient information. 
Press to input other patient information.

	5
	Record Setting
	Display recording mode and recording format.

Press to select other mode and format.

	6
	Alarm Area
	Display text alarm information, including system failure about patient cable/Print head/Paper, lead off, AC interference, EMG interference, baseline wander, and data overflow, etc.

	7
	System Status
	Indicate system status, for example, mute, recording, network, SD card inside, USB connecting, battery, etc.

	8
	Time
	Display system time.

	9
	Shortcut Keys
	Quick operation to set up parameters and execute functions.

	10
	Waveform Display
	Display real time waveforms.


· EXPLANATION
Main interface may slightly differ from the product you purchased, which will not affect your usage, please operate according to the actual functions of the product.
4.1 Introduction of Sensitivities, Filters, Paper speed
Before printing, it is necessary to do the following parameter setting:

Press[image: image66.jpg](] J)

74



to set sensitivites.
Press[image: image67.jpg]


to set low-pass filter, baseline wander filter and AC filter.
Press[image: image68.jpg]


 to set paper speed.
· EXPLANATION

· A sort of noise may degrade collected ECG signal. You can select a set of filters to optimize the displayed or printed ECG waveforms.

· While filters settings affect displayed and printed ECG waveforms, the analysis result will not be affected because algorithm receives and analyzes data filtered with fixed bandwidth.

· In order to reduce baseline wander, a baseline wander filter should be employed. To make sure the ST segment is not distorted, the AAMI standards recommend that cut-off frequency of the baseline wander filter is lower than 0.67Hz.
4.2 Recording an ECG
· EXPLANATION
· When the Pre-acquisition mode is on, the waveforms printed and uploaded are that before pressing [image: image69.jpg]


 to start, When the Pre-acquisition mode is off, the waveforms printed and uploaded are that after pressing [image: image70.jpg]


 to start.

· [Record Format] is a waveform pattern traced on the recording paper. Please see “Technical Specification” for specific recording formats.
· For ECG type of Resting 12, if “Cabrera” is selected, the "Record Mode" is automatic recording mode, and "Record Format" is a fixed format.
· [Synchronous Printing] and [Real-time Printing] take effect only when waveforms are printed in more than one columns. By selecting [Synchronous Printing], the starting times of all waveforms are the same; by selecting [Real-time Printing], the starting times of different columns are different, and the starting time of one column is sequential to the ending time of the previous column. Set them in [Print Data Type] of [Print Setting].
· When the [Print Grid] of [Print Setting] is on, it will print grids on the paper without grid. It is suggested to select “on” when using paper without grid.
· CAUTION
When copying or moving the files, it is not allowed to insert or unplug the U disk or SD card; otherwise, it may cause abnormality of the ECG machine.
4.2.1 Main operation process to record ECG


















When the waveforms recording is about to finish (under any recording mode except manual mode), the system will carry on a resting ECG analysis automatically. Please refer to Automatic Recording Mode.
4.2.2 Arrhythmia Detection
If [Arrhythmia Mode] in [ECG Setting] is set to enable, when arrhythmia is detected, rhythm waveforms can manually be printed as following:









Waveforms length is based on [Rhythm Time], please refer to Rhythm Mode.
· EXPLANATION
· Arrhythmia can only be detected in automatic mode, economic mode and cycle mode.
· Please refer to Introduction to Sensitivity, Filter and Paper Speed to set these parameters before printing.
· Please refer to Troubleshooting to understand how to avoid abnormal waveform recording, for example interference.
· Caution 
· After the heart rate and waveforms become stable, you can print out the ECG waveforms together with a resting ECG analysis.
· When the patient wears a pacemaker, the low-pass filter should be set to >150 Hz.
· In sequence to prevent the pacemaker from being detected repeatedly or omitted, you should set the detection sensitivity according to the actual clinical need. Sensitivity can be set in [Pacemaker Detection] of [ECG Setting].
· This ECG machine can detect the lead connecting status continuously, and if any leadoff is detected, the corresponding lead code will be displayed in the Alarm Area on the main interface, accompanied by sound alarm. When “Lead off” continues, please check carefully the connections from skin to the ECG machine (including electrodes, lead wires and acquisition module). Alarm will disappear when connections become reliable.
4.3 Introduction of Recording Mode
· EXPLANATION
· Waveform length to be printed and uploaded is set in [Wave Sample Time] of [ECG Setting].
· If [Auto Upload] in [ECG Setting] is enabled, the ECG machine will automatically upload the waveforms and the reports after printing the waveforms.
· If [Auto-save] in [ECG Setting] is enabled, the ECG machine will automatically store the waveforms and reports after printing the waveforms.
· If the waveform and the report need to be stored in a specified memory, you shall select the desired storage location in [Default Memory] of [System Setting], including three options: the ECG machine itself, U disk and SD card.
4.3.1 Automatic Recording Mode
In automatic mode, the ECG machine can automatically print the waveforms and reports according to [Report Format] in [Print Setting].




“Report Format”






The ECG machine has the function of resting ECG analysis and can output the measurement data, median Beat and analysis result, etc.

Simple report includes patient information, simple measurement data and Minnesota code;

Detail report includes patient information, simple measurement data, Minnesota code and detailed measurement data;

Median Beat Report includes patient information, simple measurement data, Minnesota code, Median Beat waveforms and Rhythm waveforms.

In manual report mode, the analysis reports include analysis report (I), analysis report (II) and analysis report (III), please refer to Analysis Report Mode for details.

When [Analysis Output] of [Print Setting] is enabled, the above-mentioned reports will include the analysis result.
After printing waveforms and report, press [image: image71.jpg]


 to duplicate waveforms and report of the previous patient.
· EXPLANATION
· Function of resting ECG analysis analyzes just the waveforms in the latest 10 seconds.
· If patient age is not inputted, the ECG machine will assume patient as adult during analyzing.
· WARNING
· In the case for some special populations (such as pregnant women, the user of vascular drugs, etc.) or mixed by obvious interference in the recording process, the analysis result may be inaccurate. Therefore the final conclusion should be drawn by a doctor, based on analysis result, the clinical characterization of patients and other diagnostic test results.
· If there is too much AC and EMG interference, the identification of P wave and Q wave is not reliable sometimes; if there is baseline wander, the identification of ST segment and T wave is not reliable sometimes.
· If the ending points of S wave and T wave are winding and not clear, it might cause measurement error.
· If R wave is undetected because of low voltage for QRS complex, it might cause some deviations in heart rate measurement.
· If QRS complex has low voltage, the electrical axis measurement and the identification of QRS dividing point can be unreliable.
· Occasionally, the frequent (repetitive) ventricular premature beat might be detected as the median beat.
· When several kinds of arrhythmia occur simultaneously, the identification of P wave might be difficult, and the relative parameters might be unreliable.
4.3.2 Manual Recording Mode
In manual recording mode, press [image: image72.jpg]


 to begin or stop printing. You can switch by pressing [image: image73.jpg]((<oad)) [te2a))




 from one lead group to another at any time to control the printing time of every lead group. 
In manual recording mode, the ECG machine will not analyze or measure the acquired waveforms.

4.3.3 Economic Recording Mode
In economic mode, all the waveforms and analysis result are designed to be printed on A4 size thermal paper, so that recording can save the paper costs.
· EXPLANATION
· In economic mode, only “Simple Report” can be printed, “Detail Report” and “Median Beat Report” are not supported. Economic mode is not applicable to 3-channel and 6-channel ECG machine.
· To save paper, only waveforms of 10 seconds can be printed.
4.3.4 Cycle Recording Mode
In cycle recording mode, the system will count down according to [Cycle Time], record the ECG waveforms automatically and periodically according to [Cycle Interval], until the end of time. Please refer to Print Setting.
· EXPLANATION
After being set, the cycle recording mode will not take effect immediately until pressing[image: image74.jpg]


.
4.3.5 Trigger Mode
In trigger mode, when there is an arrhythmia, the ECG machine can automatically detect it and trigger the printing of waveforms and analysis report. After printing, the ECG monitor will wait until you confirm whether to continue the monitoring or not.
· EXPLANATION
After being set, the trigger mode will not take effect immediately until pressing[image: image75.jpg]


.
4.3.6 Upload Mode
In upload mode, select [Upload] or press[image: image76.jpg]


, the ECG machine will enter into analysis. After analysis, the ECG machine will upload the waveforms and reports to ECG management system. During the process, no printing happens.
4.4 Advanced Mode
4.4.1 Rhythm Mode
Operate as follows to enter rhythm mode:










When entering rhythm interface, the ECG machine begins collecting waveforms of the Rhythm Lead.  You can select single-rhythm and three-rhythm pattern. In single-rhythm pattern, only one lead is selected as rhythm lead and as long as 300s waveforms will be collected and analyzed. In three-rhythm pattern, waveforms of three leads will be collected and analyzed, as long as 100s for each lead, totally 300s. After collecting waveforms, the ECG machine will automatically analyze the waveforms and enter report interface.
In rhythm report interface, you can [Print], [Save], [Upload] and press [Pageup] and [Pagedown] to review more information. 

Please refer to ECG Setting to set [Rhythm Mode], [Rhythm Time] and [Rhythm Lead].
Rhythm function may slightly differ from the product you purchased, for example some products do not support three-rhythm pattern, please operate according to the actual functions of the product.
· EXPLANATION
During the waveform collecting process, when the collecting time has been over 8 seconds, you can press [RR Interval] to enter report interface manually.
4.4.2 Analysis Report Mode
Operate as follows to enter analysis report mode:


[image: image77.png]ENs==]





Report (1) includes simple measurement data, minnesota code, median beat waveforms, analysis result and rhythm waveforms.

Report (2) includes detail measurement data.
Report (3) includes waveforms of all the leads.
Analysis report interface is as following:
Analysis report interface is as following:


[image: image78.png]@. 066 | 141017000, , , | 34 6.‘5 @_q:

Heart rate 66 bpr [Minng it (» o 7
interval 164 m: T
QRS duration 91 ms Lo 1
1QTC interval - 4341455 m: U
JORSIT axis| | - 13/54/45°
RV5/SV1 voltage - 1.92/14: !
I
— ++++++ Analyse Result ++++++ L L \
8110 Sinus rhythm
-~ ** Normal ECG **
5 1





[image: image79.png]141017000, , , 3%4
port(

Il m W [avR | aVvL [avF | VI | V2 | V3 | V4 | V5 | Ve
P amplitude w100 | 91 100 | 57 | 41 | 80 | 89 | 84 | 75 | 80 | 68
Pamplitude w
T amplitude V| 355 | 488 | 149 | -422 | 116 | 311 | -201 | 238 | 463 | 674 | 807 | 628
Q duration ms| 12 | 18 | 15 16 19 | 18
Q amplitude W -61 | -130 | -68 98 190 | -247
R dutation ms| 55 | 67 | 66 | 16 | 28 | 64 | 22 | 42 | 35 | 51 | 41 | 48
R amplitude V| 940 | 1422 697 | 96 | 229 |1025| 383 | 577 | 525 | 1144|1924 | 1704
S duration ms 62 | 38 59 | 62 | 52 | 24
S amplitude w 1181 -256 -1440 -1846 |-1238 | -543
Rdutation ms
Riamplitude w
S'duration ms
Samplitude w
STJ W[ 27 | 52 | 25 |38 | 2 | 38 |34 | -0 | 43 | 18 | 34 | 82
Q duration(Equ) ms
Tampltude w
Tampltude(mod) V| 330 | 45 | 145 | -413 | 98 | 273 | -244 | 64 | 257 | 514 | 7113 | 532
VAT ms| 35 | 40 | 42 | 14 | 25 | 42 | 12 | 29 | 25 | 40 | 37 | 35
QRS area 40ms W[ 17 | 39 | 22 | 28 | -2 | 30 | -34 | -30 | -14 | 15 | 32 | 39
STMD W 11 | 20 | 9 |16 | 2 | 16 | 18 | 128 | 144 | 107 | 43 | -16

2 & = = =





[image: image80.png]S R R

(:5) @ 066 | 141017000, , ,
Resting 12 AC:50 Hz0.35~100Hz  25mm/s 10 mmim
I !

1
[

1

i

I
L

150 SNER( SN DAERI SRS UARR! URE AER{ SNER.

.





In the above three interfaces, you can [Upload], [Page], [Save], and [Print]. In analysis report (1), [Analysis Result] can be manually edited by the user. Please refer to List of Interpretation codes and Corresponding Description for the details of analysis result.
4.4.3 Freeze Mode
Operate as follows to enter freeze mode:







You can freeze waveforms for 300s. 
In the freezing interface, waveforms of different pages can be browsed through Pageup/Pagedown and step can be adjusted by the user.
In the freezing interface, press “Print” to print the waveforms.
Press “Report”, and then choose the type of the report, you can print the selected report.
In the freeze mode, split-screen display of 12 lead waveforms is same as that in the main interface.

Please refer to Split-screen display for the operation.
Chapter 5 Setting System Parameters
In the main interface, press [Menu], or press[image: image81.png]


, to enter main menu. 
[image: image82.png]() ECG Setting

Lead Mode Standard Leads
Low-pass Filter 100Hz
Baseline Wader Filter 0.35Hz

AC Filter 50 Hz

Rhythm Mode Single-rhythm

Rhythm Lead 1 I

Rhythm Lead 2 V2
Rhythm Lead 3 V5
Rhythm Time 60 S

Pacemaker Detection Disable

Press Exitto return previous menu

Print
Setting
Display Patient
Sefting Information
System Data
Sefting  Management
A
%
Order Remote
Management  Report
i .E
I
e}
Factory
Maintain

Save | -9 Exit





· EXPLANATION

· In the process of parameter settings, press [image: image83.jpg]


and return to the previous page, step by step, until return to the main interface.
· Press “save” before exiting from setting interface to avoid loss of setting because of sudden power lost.
5.1 ECG Setting
Enter [ECG Setting] to set parameters about Electrocardiograph. See following table:
	Name
	Value
	Default
	Description

	Lead Mode
	Standard Leads, Cabrera
	Standard Leads
	Select one option for lead mode.

	Low-pass Filter
	25 Hz, 35 Hz, 75 Hz, 100 Hz, 150 Hz, 250 Hz
	100 Hz
	Select one option for low-pass filter, baseline wander filter and AC filter.

	Baseline Wander Filter
	0.01 Hz, 0.02 Hz, 0.05 Hz, 0.35 Hz, 0.5 Hz, 0.8 Hz
	0.35 Hz
	

	AC Filter
	OFF, 50 Hz, 60 Hz
	50 Hz
	

	Rhythm Mode
	Single-rhythm, Three-rhythm
	Single-rhythm
	Select one option for rhythm mode.

	Rhythm Lead 1
	Ⅰ, Ⅱ, Ⅲ, aVR, aVL, aVF, V1, V2, V3, V4, V5, V6
	Ⅱ
	Select lead as rhythm lead.

	Rhythm Lead 2
	Ⅰ, Ⅱ, Ⅲ, aVR, aVL, aVF, V1, V2, V3, V4, V5, V6
	V1
	

	Rhythm Lead 3
	Ⅰ, Ⅱ, Ⅲ, aVR, aVL, aVF, V1, V2, V3, V4, V5, V6
	V5
	

	Rhythm Time
	30 s～300 s
	60 s
	Select one option for rhythm recording time.

	Pacemaker Detection
	Disable, Weak, Normal, Enhance
	Disable
	Set the pacemaker detection sensitivity according to patient’s pacemaker status.

	Arrhythmia Mode
	Disable, Enable
	Disable
	Set the arrhythmia detection mode.

	Wave Sample Time
	10 s～24 s
	10s
	Select one option for waveforms sampling time.

	Pre-acquisition
	Disable, Enable
	Disable
	Set the pre-acquisition mode.

If enabled, it can print out pervious waveforms.

	Auto-save 
	Disable, Enable
	Enable
	Set whether to automatically save the report.

	Data Format
	ECG, XML, JPEG, DICOM
	ECG
	Select one option for data format.

	Auto Upload
	Disable, Enable
	Disable
	Set whether to automatically upload the waveforms and report after printing.

	QTC Formula
	Bazett, Fridercia, Framingham, Hodges
	Bazett
	Select one option for QTC formula.

	Examination Type
	Normal, Physical Examination
	Normal
	Select one option for examination type.


· CAUTION

In sequence to prevent the pacemaker from being detected repeatedly or omitted, you should set the detection sensitivity according to the actual clinical need.
· EXPLANATION
For physical examination of large numbers of people, it is recommended to set [Examination Type] to Physical Examination.
5.2 Print Setting

Enter [Printer Setting] to set parameters about printer. See following table:

	Name
	Value
	Default
	Description

	Gray Level
	1～8
	6
	Select one option for gray level.

	Baseline Width
	1～4
	2
	Select one option for waveforms thickness.

	Paper Speed
	5 mm/s, 6.25 mm/s, 10 mm/s, 12.5 mm/s, 25 mm/s, 50 mm/s
	25 mm/s
	Select one option for paper speed.

	Report Format
	Simple Report, Detail Report, Median Beat Report, Disable
	Simple Report
	Select one option for report format.

	Analysis Output
	Enable, Disable
	Enable
	Select whether to print a resting ECG analysis automatically.

	Print Grid
	Enable, Disable
	Disable
	Select whether to print the grid on the paper.

	Printer
	Built-in Thermal Printer, Disable
	Built-in Thermal Printer
	Select whether to print using built-in printer.

	Record Mode
	Refer to the appendix Technical Specification and the equipment your purchased.
	Auto
	Select one option for report mode.

	Record Format
	Refer to the appendix Technical Specification and the equipment your purchased.
	3X4
	Select one option for report mode.

	Print Date Type
	Real Time , Synchronous 
	Real Time
	Select one option for print date type.

	Cycle Time
	1~60min 
	60 min
	Select one option for cycle time and cycle interval for cycle recording mode.

	Cycle Interval
	1~setted value of cycle time min
	1 min
	


5.3 Display Setting

Enter [Display Setting] to set parameters about display. See following table:

	Name
	Value
	Default
	Description

	Display Style
	Classic White, Classic Black
	Classic White
	Set the display style of the screen.

	Background Grid
	Enable, Disable
	Enable
	Set the Background gird Disable or Enable.

	Display Format
	Same Screen, Split-screen
	Same Screen
	Set the waveforms to display in the same time (same screen) or different time (different screen).

	Lead Format
	Refer to Technical Specification.
	3X4
	Select one option for lead format.

	Lead Standard
	IEC Standard, AHA Standard
	IEC Standard
	Select a lead standard to display.


5.4 Patient Information Setting

Enter [Patient Information Setting] to set patient information. See following content:

Sub-ID No., Gender, Age, Date of birth, Height, Weight, BP, Race, Pacemaker, Medication, Accession No., Ref-physician, Technician, Physician, Ref-department., Room No., Urgent, Hospital No., User Defined.

· EXPLANATION

Age and D.O.B can’t be selected at the same time.
5.5 System Setting

Enter [System Setting] to set system parameters. See following table:

	Name
	Value
	Default
	Description

	Demo Mode
	Normal ECG, Arrhythmia ECG, Disable
	Disable
	Select a waveform type displayed in the main interface as a demo. Arrhythmia ECG is selected for demonstrating auto trigger mode.

	System Language
	Chinese, English etc.
	To be determined by the shipping country
	Set the system language.

	System Version
	Version No., Compile Time, Acquisition Module Version
	/
	Display the details of software version.

	System Time
	Current Time, Date Format, Date, Time
	/
	Display the details of time and date and to set the date format.

	Network Setting
	Cable Network, WIFI
	Cable Network
	Select one option for network setting.

	Transfer Protocol
	TCP, FTP
	TCP
	Select one option for transfer protocol.

	Cable Network
	IP Address, Subnet Mask, Default Gateway
	/
	Set the value of IP Address, Subnet Mask, and Default Gateway.

	WIFI Network (Optional)
	SSID, password and Security

DHCP: Disable, Enable

IP address, subnet mask and gateway
	/
	Set the value of SSID, password and Security.

Set DHCP to enable or disable accordingly. 

If DHCP is disabled, set the value of IP address, subnet mask and gateway.

	3G Network
	/
	/
	/

	Server Setting
	IP Address, Port
	/
	Input the value of IP Address, port of server.

	FTP Server
	User, Password, Path
	/
	Input the value of user, password and path of the FTP server.

	Silent Mode
	Disable, Enable
	Disable
	Set the silent mode Disable or Enable.

If silent mode is enabled, all sound, including alarm tone and key tone will be off.

	QRS Tone
	0～10
	6
	Select one option for three sound tones. If all three options are zero, a silence icon will be displayed in the main interface.

	Alarm Tone
	
	
	

	Key Tone
	
	
	

	Default Memory 
	Internal Memory, USB Flash Drive, SD card
	Internal Memory
	Select the default memory way for the saved file.

	Memory Format
	Internal Memory Formatting, SD Card Formatting
	Internal Memory Formatting
	Format the specified memory. The files cannot be recovered after formatting.

	Standby Time
	None, 5 min, 10 min, 30 min, 1 h, 2 h
	Disabled
	Set the standby time.

	Auto Power-Off
	None, 30 min, 1 h, 2 h, 3 h
	Disabled
	Set the auto power off time.

	System Password 
	Disable, Enable
	Disable
	Set the system password disable or enable.

	Password Setting
	0～9999
	1234
	Set the value of password when the system password is enabled.

	General Setting
	1～10
	1
	Select one option, and then set up according to your habit, all your setup will be stored in this option to facilitate your next use. Different doctor or different check up can occupy different options.

	Import Setting
	Import from USB flash disk, 

Import from SD Card
	/
	Import files from the specified memory.

	Export Setting
	Export from USB flash disk, 

Export from SD Card
	/
	Export files to the specified memory.

	Touch Calibration
	/
	/
	Make calibration to the touch screen.

	Factory Default
	/
	/
	Restore to factory default.

	Hospital
	/
	/
	Input the name of hospital.

	Device No.
	/
	/
	Input the number of this ECG machine.


· caution

Demo mode is designed for representation only. Do not use this mode in clinical analysis, for demo waveforms may be mistaken as that of patient and misdiagnose may happen.
5.6 Remote Report
Press [Setting] to input the “Start Time” and “End Time”, and then press [Load], the ECG files of that duration can be downloaded into the ECG machine and displayed on the interface below:

[image: image84.png]Select All
Shortcut Button

ECGFiles

File Atfributes Area

47 Remote Report

Date

Display Area

2014/01/13,14:59:29 140113025
2014/01/13,14:5842 140113024
2014/01/13,14:09:26 140113023
2014/01/13,14:0752 140113022
2014/01/13,12:5857 140113021
2014/01/13,12:5746 140113020
2014/01/13,12:5459 140113019
2014/01/13,12:54.07 140113018
2014/01/13,12:53:34 140113017

Neme

Gender  Age

@B | Peoeve | Psceoown| open [ sewng |

Page: 1/1  Totalfiles: 9
1

Type

Resting 12
Resting 12
Resting 12
Resting 12
Resting 12
Resting 12
Resting 12
Resting 12

Resting 12

Indication Area





In the above interface, you can [Page Up], [Page Down], [Open] and etc. The operations in this interface are the same as those in Data Management.
5.7 Factory Maintain

Only the authorized service engineer can set up [Factory Maintain], please contact with our customer service if necessary.
--Blank Page--
Chapter 6 Data Management

Press [Data Management] in the main menu interface to enter the data management interface. Select one source of storage medium from Local (ECG machine itself), U disk and SD card, the ECG files will be uploaded.
[image: image85.jpg]Viv' Data Management X ]
Date D Name Gender | Age Type
2011014170210 141014000 Rhythm
2014110/14,16:1131 141014012 Rhythm
201411014,16:0706 141014011 Rhythm
2014110/14,1539:10 141014010 Resting 12
2014110/14,153636 141014009 Resting 12
2014110/14,1536:04 141014008 Rhythm
2014110/14,153654 141014007 Resting 12
Open Upload PageUp || PageDown || Operate v || ECG || Local ~
Page: 1/1 201950 V|




	No.
	Name
	Description

	1
	Select-All Shortcut Button
	Check to select all the ECG files in the current page.

	2
	File Attributes Area
	Select any file attribute, and sort files by pressing [image: image86.jpg]


on[image: image87.jpg]


.

	3
	ECG Files Display Area
	Display basic patient information of all the ECG files.

	4
	Open
	Open one patient’s ECG file.

	5
	Upload
	Upload selected ECG files to ECG management system or server.

	6 and 7
	PageUp/PageDown
	Browse ECG files in previous or next page.

	8
	Operate
	For “Select all”, “refresh”, “copy”, “move”, “delete”, and “search” the ECG files.

Search ECG files according to ID number, name, age, time and symptom.

	9
	ECG File Type
	Select one option for ECG file’s format: including ECG, XML, JPEG, DICOM and PDF.

	10
	Select Storage Medium
	Select one option for the storage medium, including local, USB flash drive, and SD card.

	11
	Indication Area
	Indicate pages of ECG files and internal memory.


6.1 Open an ECG File

Check an ECG file, and then press [Open] or press the shortcut left or right direction key of [image: image88.jpg]


, the file will be opened.

· EXPLANATION

When you select more than one file to open, the default file is the first selected file.
6.2 Edit an ECG File

After opening the ECG file, you can edit the patient information and analysis result, can also upload, save and print the file. You can refer to Analysis Report Mode to know the content and function of the ECG file.

6.3 Delete ECG Files
Press “delete” in the [Operation] to delete selected file or files.
· caution

Deleted files cannot be recovered. Please use this option cautiously.
6.4 Copy and Move ECG files

Select one or more ECG files, operate as following to transfer the ECG files.
[image: image89.jpg][Data Management]

v

Select one or several files

[Operate]

Choose the medium copy or move to

v
ECG files copy or move to the selected medium




· EXPLANATION

· Files can be copied or moved between the local ECG machine and external storage medium. ECG files on the local machine will be deleted if the user moves them.
· After selecting files and a path, the files will be copied to the selected path. When copying files to SD card and USB disk, the system will create a new directory in the SD card and USB disk to store the selected files, for example “ecg database” etc.

· When there is no enough memory, the system will indicate that. You should select new memory to make files copying or moving successfully.

· Please regularly clear data in a storage medium; otherwise, the speed of the machine will be slowed down.

· caution
· When copying or moving the files, the continuity of power supply must be assured, or the files may be missing.

· When copying or moving the files, it is not allowed to insert or unplug the U disk or SD card; otherwise, it may cause abnormality of the ECG machine.
Chapter 7 Maintenance

7.1 Main Unit
· CAUTION

· Gently disconnect the acquisition module and power cable without forcibly pulling the lead wires.

· Clean the ECG machine as well as the accessories periodically and cover it from dust. 

· Store the unit in a dry and cool environment and avoid excessive shocking and vibration.
7.2 Patient Cable
· CAUTION

· The lead wires must be periodically checked for good connection. Damage may cause abnormal ECG waveforms at some or all leads.

· The user should avoid twisting the patient cable, or the life time will be shortened.
7.3 Cleaning and Disinfection
Before cleaning, power off the ECG machine and disconnect it from the AC power.

Do cleaning first before disinfection.
The process to clean and disinfect the ECG machine, cables, lead wires and reusable electrodes are as follows:
1) Use a clean soft cloth absorbing an amount of cleanser or disinfectant to wipe the surface carefully and avoid touching connectors of the ECG machine and accessories.
2) When necessary, wipe the superfluous cleanser or disinfectant with dry cloth.
3) Place in the ventilated and cool environment to dry the ECG machine and accessories.
Sterilization operation for this ECG machine and accessories is not recommended, unless the manual of the accessories has requirement.
· CAUTION
· While cleaning and disinfection, do not splash liquid into the ECG machine and the accessories.
· Disinfections may cause damage to the ECG machine or accessories to a certain degree. It is suggested that only when necessary, disinfect the ECG machine and accessories.
· Neutral cleanser or disinfectant is recommended.
7.4 Recording Paper
· CAUTION

· To ensure good ECG recording, please use suitable thermal recording paper for the ECG machine. Incorrect recording paper can damage printer head and cause problems such as blurring trace and incorrect paper running. Pay attention to the following comments on recording paper.
· Never use recording paper coated with wax for the ECG machine. It may cause serious problem to the printer head.
· When exposed to high temperature, high humidity and direct sunlight, the recording paper will deteriorate. It is therefore required to store the thermal recording paper in a dry and cool environment.
· When exposed to fluorescent light for long time, the recording paper will deteriorate.

· When stored with polyvinyl chloride (PVC), the recording paper will deteriorate.
· If the thermal recording paper is stored overlapping for a long time, the printing impression will leave traces in other pages, which will cause mislead readings.

· Use suitable size recording paper for the ECG machine. Or, it may cause damage to the printer head and Silicon rubber shaft.
7.5 Battery
The ECG machine is equipped with a built-in rechargeable battery to assure continuous operation when AC power is unavailable. Charging, capacity indication and replacement of the battery are described below:
· Charging

The ECG machine is designed with a charger and protector for the battery.

· Please turn off the machine first before charging the battery.

· The battery-charging indicator on the operation keyboard will become green when the battery is charged completely.

· Discharge and charge the battery at least once every three months (discharge the battery until the machine turns off automatically, and then fully charge the battery).

· Capacity indication
When the unit is powered by battery, there will be a symbol of battery capacity indication displayed on the LCD.

For example:

[image: image90.png]


 Full battery capacity, it can work continuously for about 3 hours.
[image: image91.png]


 Battery capacity is sufficient.

[image: image92.png]


 Insufficient battery capacity, charging is required.

[image: image93.png]


 Battery capacity is going to running out, immediate charging is demanded.

[image: image94.png]


 Battery capacity has already run out and blackout may happen at once, immediate charging is demanded.

· Battery replacement

The battery should be replaced by the professionals according to the following procedures.
1. Power off the ECG machine and disconnect the AC power cable.
2. Flip over the ECG machine and disassemble the battery back cover based on the instruction on the back cover.
3. Disconnect the battery plug and take out the battery.
4. Replace the existing battery with a new one. Pay attention to polarity and connection.
5. Install the back cover.
[image: image95.png]



· EXPLANATION

Refer to Foreword for other warning information of battery.

7.6 Silicon Rubber Shaft for Printing
The silicon rubber shaft shall be kept clean, smooth and free from dirt. Otherwise, the ECG machine may print out unsatisfied ECG trace. To clean dirt from the shaft, wipe the shaft with soft cotton moistened with medical alcohol and at the same time rotate shaft until it is clear enough.

7.7 Thermal Printing Head
Residue and dirt on the thermal printer head could affect the clarity of recorded ECG waveforms. To clean the thermal printer head, open the paper cover and clean the printer head with soft cotton moistened with medical alcohol. It is not permitted to clean the printer head with a sharp object, which can cause permanent damage to the printer head. Thermal printer head maintenance should be done at least once a month.

Chapter 8 Troubleshooting
8.1 Lead Fault

1. Data saturation or overflow happens.

Solution:

Ensure that all leads are in good contact, and wait for half a minute or the waveforms on the screen are stable, and then start printing.
2. Straight line is printed in some leads.
Solution：
1) Check if the metal piece of limb electrode contacts with the body properly; if not, adjust the position of the limb electrode, and adjust the tightness if necessary.
2) Check if the limb electrodes and chest electrodes are oxidized or faded, and clean the accessories or replace with new ones. Oxidation and aging cause conductive deterioration of the electrodes, resulting in poor signal transmission.
3) Treat the skin of the patient with alcohol; because dry skin causes skin resistance to become larger, which will impact the signal collecting.
4) Please clean joints of lead wires, suction ball and limb clip, reinstall and tighten all joints. After long-term use, joints will have dirt or become loose, resulting in poor signal transmission.
5) Check if the appearance of the lead wires has obvious fracture; if yes, replace with new lead wire. If not, connect a proper lead wire to the device. If the waveforms are stable, the lead wires have problems and have to be replaced.
6) If there is no lead wires available, check if the lead wires conduct with a multimeter. First check if the inner conductors of the lead wires are conductive. Generally speaking, the acceptable resistance shall be about 10 kΩ. Then check if there is a short circuit between the outer shield and inner conductor. The resistance shall be infinity. If the lead wires have a problem, please contact our customer service to replace new lead wires.
3. Lead fault can also be caused by failure in signal communication. Please exclude other causes for lead fault problems first, and then contact our service department if necessary.

8.2 Printer Failure

1. Unclear printing.
Solution:
1) Whenever a printer fault occurs, such as poor or incorrect ECG recording, you may try to clean the thermal printer head with soft cotton dipped with medical alcohol.

2) If the quality of the thermal paper is poor or the paper is not used for a long time, which may result in reduced performance of thermal layers. Please replace with provided or specified recording paper.

3) If the above methods are not applicable, guide the user to test the print head and check if the print head has breaking point; if yes, contact the customer serviceto replace the thermal print head.

2. Upper half or lower half is blank.
Solution：
Check if the bearing on both ends of the rubber shaft of the paper compartment cover is worn, and replace with new bearing if yes.
3. All paper or most part is blank.
Solution:

1) Make sure that the thermal recording paper is not installed backwards.

2) Check if the print head is stuck by dirt (such as adhesive tape); this often occurs when new print paper is replaced.

8.3 Keyboard Failure

If key failure happens, the professional maintenance staff can locate the failure by "Key Test" in [Factory Maintain]. 
Keyboard failure is probably caused by loose connection between the keyboard panel and the System Control Board due to shocking or vibration during transportation. If necessary, a professional service engineer will disassemble the ECG machine and reconnect the inner cable. At the same time, check keyboard, if key may be stuck for aging.
8.4 Indication of Lead Off

This ECG machine can detect the lead connecting status continuously. When the leads are not well connected to the main unit, it means that the signals cannot be transferred correctly, thus there is “lead off:*” indication, accompanied by voice alarm. The symbol “*” represents the fault lead, the waveforms of which will display as a straight line. Please check carefully whether the connection among the related electrodes, human body, patient cable and the main unit remains well.
8.5 AC Interference

Apparent and regular trembling of ECG waveforms in the process of recording due to AC interference is shown as below.
[image: image96.png]



Causes of baseline wander are varied, please do following checks one by one：
1. Make sure the ECG machine is properly grounded according to the instructions.

2. Make sure the patient cable and electrodes are properly connected.

3. Make sure the electrodes and the patient skin have been covered with ECG gel.

4. Make sure the exam bed is properly grounded.

5. Make sure the patient is not in touch with the wall or the metal part of the bed.

6. Make sure the patient is not in touch with anybody else.

7. There shall be no large power electric equipment (such as X ray machine, ultrasound scanner etc.) operating nearby.

8. The patient shall not be wearing some jewelry like diamond.

· CAUTION

Set AC filter to ON if AC interference still exists after the above checks are completed.
8.6 EMG Interference

Irregular trembling of ECG waveforms due to EMG interference is shown as below.

[image: image97.png]



Causes of baseline wander are varied, please do following checks one by one：
1. Make sure that the exam room is comfortable for examination.

2. Soothe the patient from irritation or excitement.

3. Make sure the exam bed is in suitable size.

4. Never have talks with the patient during ECG recording.

5. Make sure the limb electrode is too tight to make the patient uncomfortable.
· CAUTION

Set EMG filter to lower value if EMG interference still exists after the above checks are completed, and keep in mind that the recorded ECG waveforms, particularly R wave, will be somewhat attenuated.

8.7 Baseline Wander

Irregular movement of ECG baseline due to baseline drift is shown as below:
[image: image98.png]



Causes of baseline wander are varied, please do following checks one by one：
1. Make sure the electrodes are in good contact with skin.
2. Make sure the connection between the patient cable and electrodes is normal.
3. Make sure the electrodes are clean and patient skin contacted with the electrodes has been well pretreated.
4. Make sure the electrodes and skin are covered with ECG gel.
5. Keep the patient silent and motionless, and keep the patient from hyperventilation.
6. Used electrodes shall not be utilized with new ones in patient examination.
· CAUTION

If the problem still can’t be cleared, please turn up the baseline wander filter, keep in mind that the recorded ECG waveforms, particularly T wave and ST segment, will be somewhat distorted.
8.8 The ECG machine cannot be turned on

1. AC power is not working properly and the battery is exhausted.

Solution:

First check if the power outlet connects properly, if the power line and the machine connect properly, then check if local AC voltage is normal. If everything is ok, check if the fuse is good. If everything is normal after above examination, return the machine to the manufacturer for repair, for it may be damaged.
2. After turning on, the machine turns off automatically after a few minutes.

Solution:

While working, if the screen displays the battery power, the machine is using the battery, but the battery power is insufficient, resulting in automatic shutdown. Please supply the machine with AC power, or charge the battery before working. If the user is using AC power and the machine still turns off automatically, please check according to step1.
8.9 Paper Feeding Failure

1. Press[image: image99.jpg](&)



, but paper is not fed.

Solution:

Check if the keyboard has been damaged and replace damaged keyboard.

2. Press[image: image100.jpg](&)



, the paper is not fed, and there is abnormal sound from recorder.

Solution:

First check if the recording paper is installed properly, and if the gear on the cover of paper magazine is in good condition.

3. Paper feeding isn’t smooth, paper is stuck, or waveforms are compressed.
Solution:

First check if the thermal paper complies with the standard, then if the paper is installed properly, or replace a new roll of paper. Finally, replace the paper shaft.
8.10 Battery is quickly charged and discharged

If the battery is often not fully charged, the performance will be deteriorated.

Solution:

It is recommended to charge the battery continuously until the battery is fully charged and activated for the first two times. Supply the machine with AC power as far as possible.
8.11 Wrong Analysis Result

For the case for some special populations (such as pregnant women, the user of vascular drugs, etc.) or mixed by obvious interference in the recording process, the analysis result of the resting ECG analysis of this ECG machine may be inaccurate. The possible reason may be as followings:
1. Poor contact between electrode and patient skin, caused by improper skin treatment and incorrect connection.

2. The patient has relatively large movement in the recording process.

3. Gender or age isn’t entered;

4. If there is too much AC, EMG and breathing interference, the identification of P wave and Q    wave is not reliable sometimes; if there is baseline wander, the identification of ST segment and T wave is not reliable sometimes.

5. If QRS complex has low voltage, R-wave may be missed, and the electrical axis measurement and the identification of QRS dividing point can be unreliable. Or frequent ventricular contraction occurs or a variety of arrhythmias merge, the relevant detection parameters may be unreliable.
6. The filter settings are incorrect.
Solution:

1) Treat as Apply Electrode and wait until the waveforms are stable before reanalyzing.

2) Enter the patient gender and age correctly.

3) Exclude the interference as the methods described in AC Interference, EMG Interference and Baseline Wander before reanalyzing.
4) Reset to an appropriate filtering value.
8.12 File Uploading Failure

The most possible reason is that network settings have problems, please check the network connection and refer to Network Connection to re-set the network.
--Blank Page--

Appendix A Package and Accessories
A.1 Packing List
	Type
	Item
	Qty

	Standard
	ECG machine
	1 unit

	
	acquisition module
	1 set

	
	limb electrode
	1 set

	
	chest electrode
	1 set

	
	power cable
	1 piece

	
	power adaptor
	1 piece

	
	equipotential cable
	1 piece

	
	thermal recording paper
	1 piece

	
	paper shaft(only applicable for roll paper)
	1 piece

	
	user manual
	1 copy

	
	qualified certificate
	1 copy

	
	packing list
	1 copy

	Optional
	electrode adaptor and electrodes
	1 set

	
	SD card
	1 piece


A.2 Dimensions and Weight 
	Length × Width × Height
	324 mm×264 mm×95 mm (3-channel)

	
	257 mm×291 mm×106 mm (6-channel)

	Net Weight
	About 2.3 kg (3-channel)

	
	About 2.5 kg (6-channel)


· CAUTION

· Please open the package according to instructions on the packing box.
· Accompanying accessories and documents shall be checked according to the packing list before starting checking on the unit.
· Whenever there will be mismatch of the accompanying materials with the packing list, contact our customer service immediately.
· To ensure good performance and safe operation of the ECG machine, please use the accessories supplied by the manufacturer.
· The package box should be kept well for the regular inspection or maintenance for the machine.
Appendix B Technical Specification
B.1 Specifications
B.1.1 Main unit

	Lead
	Standard 12-lead

	Acquisition Mode
	Simultaneous 12-lead Acquisition

	Recording Format
	1×12, 1×12+1R, 3×4, 3×4+1R (3-channel)

	
	3×4,3×4+1R,3×4+3R,6×2,6×2+1R (6-channel)

	Record Mode
	Auto, Manual, Upload (3-channel)

	
	Auto, Manual, Upload, Cycle, Trigger (6-channel)

	Lead Format
	3×4, 3×4+1R, 6×2, 6×2+1R, 12×1

	Rhythm Time
	30~300s waveforms acquisition for rhythm analysis

	Measurement Parameters
	Ventricular Rate, PR Interval, QRS Time Limit, QT/QTC Interval, P/QRS/T Axis, RV5/SV1 Amplitude and RV5+SV1 Amplitude

	Filters
	AC Filter

	
	Baseline Wander Filter

	
	EMG Filter

	Input CIR Current
	≤0.1 μA

	Input Impedance
	≥2.5 M(

	Time Constant
	≥3.2 s

	Frequency Response
	0.05 Hz～250 Hz

	Noise Level
	≤15 μVp-p

	Sensitivity Threshold
	≤20 μVp-p

	Sensitivity
	Auto, 1.25 mm/mV, 2.5 mm/mV, 5 mm/mV, 10 mm/mV, 20 mm/mV, 10/5 mm/mV, 20/10 mm/mV

	Standard Sensitivity
	10 mm/mV ± 2%

	Calibration Voltage
	1 mV±5 %

	Accuracy of Input Signal Reproduction
	Using the method described in 4.2.7.1 of AAMI EC11 to test the overall system error, which is within ±5%;

Using method A and D described in 4.2.7.1 of AAMI EC11 to test frequency response.
Because of sampling characteristics and the asynchronism between sample rate and signal rate of the ECG machine, digital systems may produce a noticeable modulating effect from one cycle to the next, particularly in pediatric recordings. This phenomenon, which is not physiologic, shall be clearly described in the operator's and service manuals.

	CMRR
	＞89 dB

	Patient Leak Current
	＜10 μA

	Sampling rate of signals
	8000 Hz


B.1.2 Recorder Specification

	Recorder
	Thermal Dot Matrix Word Printing System

	Recording Paper
	80 mm, roll paper or 80 mm×90 mm, z-fold paper (3-channel)

	
	110 mm×140 mm, z-fold paper (6-channel)

	Paper Speed
	(5, 6.25, 10, 12.5, 25, 50)mm/s ± 5%


B.1.3 Other Specification
	Acquisition Module
	Standard 12-lead acquisition module with defibrillation-proof

	Display on LCD
	7-inch TFT LCD screen (3-channel)

	
	8-inch TFT LCD screen (6-channel)

	Safety Classification
	IEC60601-1, Class I, Type CF

	AC Power Supply
	100 V～240 V, 50 Hz/60 Hz, 80VA (3-channel)

	
	100 V～240 V, 50 Hz/60 Hz,80VA (6-channel)

	DC Power Supply
	Rechargeable lithium battery, 14.8 V/ 2200mAh.

In environment temperature ranging from 20 ℃ to 30 ℃ and with the machine turning off, the charging time is not more than 2 hours to charge the battery to 90%.

	
	In environment temperature ranging from 20 ℃ to 30 ℃, the continuous working time is not less than 3 hours while the ECG device is continuously printing.


B.2 Environment Requirements

	1
	Transportation

	
	Environment Temperature
	-20 ℃～+55 ℃

	
	Relative Humidity
	≤95 % (No condensation)

	
	Air Pressure
	70 kPa～106 kPa

	
	Transportation: avoid direct sunshine and rain.

	

	2
	 Storage

	
	Environment Temperature
	-20 ℃～+55 ℃

	
	Relative Humidity
	≤95 %  (No condensation)


	
	Air Pressure
	70 kPa～106 kPa

	
	The packed ECG should be stored in the well-ventilated room without corrosive gases.

	

	3
	Using

	
	Environment temperature
	+5 ℃～+40 ℃

	
	Relative humidity
	≤95 % (No condensation)

	
	Air pressure
	70 kPa～106 kPa


--Blank Page--
Appendix C Working principle and block diagram
C.1 Power supply subsystem

(1) Working principle:
By ways of AC to DC high frequency power switching technique, the switching power supply output 20V DC voltage. This voltage works to charge the built-in rechargeable battery at constant voltage and limited current, and at the same time, applies to the power switchover circuitry together with the battery output. If the switchover circuitry is turned on, several stable power outputs will be generated through various switching power technique, including main power supply of +5V and +24 V（+12 V） by the switching power stabilizer, +3.3 V, +1.8 V and -5 V by the power supply transformation.

+5V, +3.3 V, +1.8 V and -5 V supply power for the control system circuitry with the equipotential as reference with 750mA normal load, 3A output current capacity and 3.75A output current limit for short circuit protection.

The +24 V（+12 V） voltage supplies power for the paper driving motor and the thermal printer. The motor is driven by means of width modulation and wave chopping technique in sequence to improve power efficiency. It has about 500mA as normal load, 850mA as output current capacity and 1.2A output current limit for short circuit protection. The self-excitation power switching circuitry transforms the output of the switchover circuitry to several voltages for analog circuitry. The +5V voltage needed by isolated digital circuitry is the direct stabilized output of the switching power supply, of which the normal load is about 150mA and the current capacity is 300mA. The +8V and -8V voltage for the isolated analog circuitry is the un-stabilized output from the switching power supply. Its normal load is about 60mA and has 100mA current capacity.
(2) Block diagram (Schematics and parts of list of this unit are only provided for qualified service center under supervision of the manufacturer.)

[image: image101.emf]Switch power 

source

DC-DC 

converter

DC-DC 

converter

DC-DC 

converter

Logic switch

Keyboard

Charger

CF-type part

Recorder

Reset

Control 

system

Battery

14.8V

AC 100-240V

50/60Hz

±8V

+5V

+24V

+5V

-5V

+3.3V/+1.8V

(+12V)


C.2 Acquisition module

(1) Working principle
Acquisition Module is connected to the main unit by ECG port which provides power supply and communication interface. When connecting, the analog circuit and control circuit of acquisition module are powered through the isolated DC/DC switching output. The analog circuit is composed of input protection circuit, anti-aliasing low-pass filter and ECG analog front chips. The electrodes acquire the millivolt electrical signals from human body, which will be converted to digital signals by ECG analog front chips first, then transferred to the processor that controls the converter in the ECG analog circuit, ECG data processing and operation keys on acquisition module. The sampling rate of signals is 8000Hz with 250Hz as bandwidth (-3dB), which meets the standard of AHA and CSE (sampling rate no less than 500Hz). After the signals are processed and filtered, they are sent back to the main unit through optically coupled isolation interface.
(2) Block diagram：

[image: image102.emf]M

a

i

n

 

U

n

i

t

DC/DC

Isolated power

Signal isolation

L

e

a

d

 

w

i

r

e

Analog circuit 

section

Microprocessor

Acquisition 

module keys


(3) Leads of acquisition module:
	Lead nomenclature
	Definition
	Name of lead

	Ⅰ
Ⅱ
Ⅲ
	Ⅰ=L-R

Ⅱ=F-R

Ⅲ=F-L
	Biopolar limb leads

(Einthoven)

	aVR

aVL

aVF
	aVR=R-(L+F)/2

aVL=L-(R+F)/2

aVF=F-(L+R)/2
	Augmented leads

(Goldberger)

	V1

V2

V3

V4

V5

V6
	V1=C1-(L+R+F)/3

V2=C2-(L+R+F)/3

V3=C3-(L+R+F)/3

V4=C4-(L+R+F)/3

V5=C5-(L+R+F)/3

V6=C6-(L+R+F)/3
	Unpolar chest leads

(Wilson)


C.3 Control system

(1) Working principle
The control system consists of controllers for printer, keyboard, LCD screen and a CPU subsystem. Through high-speed photo-couplers, CPU subsystem receives lead signal from the Data acquisition subsystem and applies them to the printer controller after digital filtration, gain adjustment and printer driving, and then complete the ECG waveforms printing. The lead data will also be measured and interpreted by CPU. In addition to measurement and a resting ECG analysis on the printed ECG waveforms, CPU also receives interruption signals and key codes from the keyboard controller to carry out key interrupt routine. Acquisition and processing of signals for detection of leadoff, out-of-paper detection, battery capacity management, automatic power off, CRO output and EXT input are all managed by CPU. The printer controller receives instructions and data from the CPU, and works to manage the buffering area and generate control signals for the stepping motor and thermal print head to print out ECG waveforms and related information.The keyboard controller works to generate keyboard scan signals, removes key bounces when key is pressed, and sends key codes and interruption signals to CPU for further processing. LCD controller receives instruction s and data from CPU, and works to display the unit’s operation status
(2) Block diagram


[image: image103.emf]Program 

memory

Microprocessor

CRO EXT

Data memory

Real time 

clock

Analog board

LCD 

controller

Recorder

LCD Programmable 

logic unit

Interface

Keyboard


Appendix D List of Interpretation codes and Corresponding Description

	8  Arrhythmia

	8002
	Marked rhythm irregularity

	8110
	Sinus rhythm

	8102
	Sinus arrhythmia

	8108
	Marked sinus arrhythmia

	8120
	Sinus tachycardia

	8130
	Sinus bradycardia

	8200
	Atrial rhythm

	8210
	Atrial fibrillation

	82101
	Atrial fibrillation with rapid ventricular response

	82102
	Atrial fibrillation with slow ventricular response

	82103
	Atrial fibrillation with aberrant conduction, or ventricular premature complexes

	82108
	Atrial fibrillation with rapid ventricular response with aberrant conduction, or ventricular premature complexes

	82109
	Atrial fibrillation with slow ventricular response with ventricular premature complexes

	8220
	Atrial tachycardia

	8250
	Atrial flutter

	82503
	Atrial flutter with aberrant conduction or ventricular premature complexes

	82505
	Cannot rule out atrial flutter

	8300
	Junctional rhythm

	8320
	Junctional tachycardia

	8400
	Supraventricular rhythm

	8420
	Supraventricular tachycardia

	8430
	Supraventricular bradycardia

	8470
	with occasional supraventricular premature complexes

	8474
	with frequent supraventricular premature complexes

	8475
	with frequent supraventricular premature complexes in a pattern of bigeminy

	8500
	Ventricular rhythm

	8520
	Ventricular tachycardia

	8570
	with occasional ventricular premature complexes

	8574
	with frequent ventricular premature complexes

	8575
	with frequent ventricular premature complexes in a pattern of bigeminy

	8901
	Undetermined regular rhythm

	8902
	Undetermined rhythm

	8970
	with occasional ectopic premature complexes

	8974
	with frequent ectopic premature complexes

	8975
	with frequent ectopic premature complexes in a pattern of bigeminy

	8706
	Electronic atrial pacemaker

	8707
	Electronic ventricular pacemaker

	8708
	Dual chamber Electronic pacemaker

	8709
	Demand pacemaker

	8710
	Pacemaker failure

	85201
	Ventricular Fibrillation

	6  AV Conductive Defect

	611
	Possible third degree AV block

	621
	Second degree AV block, Wenckebach type

	622
	Second degree AV block, Mobitz type II

	623
	First degree AV block

	631
	Short PR interval

	641
	Type-A WPW syndrome

	642
	Type-B WPW syndrome

	643
	Atypical WPW syndrome

	644
	Intermittent WPW syndrome

	7  Intraventricular Conductive

	711
	Left bundle branch block

	712
	Incomplete left bundle branch block

	721
	Right bundle branch block, plus  RVH

	722
	Right bundle branch block

	723
	Incomplete right bundle branch block

	724
	RSR' in lead V1/V2, consistent with right ventricular conduction delay

	731
	Left anterior fascicular block

	732
	Left posterior fascicular block

	741
	Nonspecific intraventricular conduction block

	742
	Nonspecific intraventricular conduction delay

	1  Myocardial

	1113
	Cannot rule out anterior myocardial infarction, probably old

	1114
	Cannot rule out anterior myocardial infarction, age undetermined

	1121
	Possible anterior myocardial infarction, possible acute

	1122
	Possible anterior myocardial infarction, probably recent

	1123
	Possible anterior myocardial infarction, probably old

	1124
	Possible anterior myocardial infarction, age undetermined

	1131
	Anterior myocardial infarction, possible acute

	1132
	Anterior myocardial infarction, probably recent

	1133
	Anterior myocardial infarction, probably old

	1134
	Anterior myocardial infarction, age undetermined

	1213
	Cannot rule out anteroseptal myocardial infarction, probably old

	1214
	Cannot rule out anteroseptal myocardial infarction, age undetermined

	1221
	Possible anteroseptal myocardial infarction, possible acute

	1222
	Possible anteroseptal myocardial infarction, probably recent

	1223
	Possible anteroseptal myocardial infarction, probably old

	1224
	Possible anteroseptal myocardial infarction, age undetermined

	1231
	Anteroseptal myocardial infarction, possible acute

	1232
	Anteroseptal myocardial infarction, probably recent

	1233
	Anteroseptal myocardial infarction, probably old

	1234
	Anteroseptal myocardial infarction, age undetermined

	1313
	Cannot rule out anterolateral myocardial infarction, probably old

	1314
	Cannot rule out anterolateral myocardial infarction, age undetermined

	1321
	Possible anterolateral myocardial infarction, possible acute

	1322
	Possible anterolateral myocardial infarction, probably recent

	1323
	Possible anterolateral myocardial infarction, probably old

	1324
	Possible anterolateral myocardial infarction, age undetermined

	1331
	Anterolateral myocardial infarction, possible acute

	1332
	Anterolateral myocardial infarction, probably recent

	1333
	Anterolateral myocardial infarction, probably old

	1334
	Anterolateral myocardial infarction, age undetermined

	1413
	Cannot rule out septal myocardial infarction, probably old

	1414
	Cannot rule out septal myocardial infarction, age undetermined

	1421
	Possible septal myocardial infarction, possible acute

	1422
	Possible septal myocardial infarction, probably recent

	1423
	Possible septal myocardial infarction, probably old

	1424
	Possible septal myocardial infarction, age undetermined

	1431
	Septal myocardial infarction, possible acute

	1432
	Septal myocardial infarction, probably recent

	1433
	Septal myocardial infarction, probably old

	1434
	Septal myocardial infarction, age undetermined

	1513
	Cannot rule out lateral myocardial infarction, probably old

	1514
	Cannot rule out lateral myocardial infarction, age undetermined

	1521
	Possible lateral myocardial infarction, possible acute

	1522
	Possible lateral myocardial infarction, probably recent

	1523
	Possible lateral myocardial infarction, probably old

	1524
	Possible lateral myocardial infarction, age undetermined

	1531
	Lateral myocardial infarction, possible acute

	1532
	Lateral myocardial infarction, probably recent

	1533
	Lateral myocardial infarction, probably old

	1534
	Lateral myocardial infarction, age undetermined

	1613
	Cannot rule out inferior myocardial infarction, probably old

	1614
	Cannot rule out inferior myocardial infarction, age undetermined

	1621
	Possible inferior myocardial infarction, possible acute

	1622
	Possible inferior myocardial infarction, probably recent

	1623
	Possible inferior myocardial infarction, probably old

	1624
	Possible inferior myocardial infarction, age undetermined

	1631
	Inferior myocardial infarction, possible acute

	1632
	Inferior myocardial infarction, probably recent

	1633
	Inferior myocardial infarction, probably old

	1634
	Inferior myocardial infarction, age undetermined

	16132
	Cannot rule out inferior myocardial infarction with posterior extension, probably old

	16142
	Cannot rule out inferior myocardial infarction with posterior extension, age undetermined

	16212
	Possible inferior myocardial infarction with posterior extension, possible acute

	16222
	Possible inferior myocardial infarction with posterior extension, probably recent

	16232
	Possible inferior myocardial infarction with posterior extension, probably old

	16242
	Possible inferior myocardial infarction with posterior extension, age Undetermined

	16312
	Inferior myocardial infarction with posterior extension, possible acute

	16322
	Inferior myocardial infarction with posterior extension, probably recent

	16332
	Inferior myocardial infarction with posterior extension, probably old

	16342
	Inferior myocardial infarction with posterior extension, age undetermined

	1711
	posterior myocardial infarction, possible acute

	1712
	posterior myocardial infarction, age undetermined

	173
	Abnormal Q wave ? [ Lat., Inf. ]

	174
	Abnormal Q wave ? [Ant. ]

	175
	Abnormal Q wave ? [ Ant., Lat. ]

	176
	Abnormal Q wave ? [ Ant., Inf. ]

	177
	Abnormal Q wave ?

	3  Ventricular Hypertrophy and Atrial

	311
	Possible right ventricular hypertrophy

	312
	Right ventricular hypertrophy

	313
	Right ventricular hypertrophy, probably repolarization abnormality

	321
	Minimal voltage criteria for LVH

	322
	Possible left ventricular hypertrophy

	323
	left ventricular hypertrophy

	324
	Left ventricular high voltage(moderate)

	325
	Left ventricular hypertrophy, probably repolarization abnormality

	331
	Possible left atrial enlargement

	332
	Left atrial enlargement

	341
	Possible right atrial enlargement

	342
	Right atrial enlargement

	3120
	Biventricular hypertrophy

	3121
	Biventricular hypertrophy with repolarization abnormality

	2  Axis Deviation

	21
	Moderate left axis deviation

	22
	Abnormal left axis deviation

	23
	S1-S2-S3 pattern

	24
	Abnormal right axis deviation

	25
	Moderate Right axis deviation

	26
	Indeterminate axis

	2  ST-T Abnormality

	2101
	ST depression, possible digitalis effect

	2102
	Minimal ST depression

	2103
	Moderate ST depression

	2104
	Marked ST depression, possible subendocardial injury

	2105
	Marked ST depression, possible subendocardial injury or Digitals effect

	2106
	Marked ST depression, consistent with subendocardial injury

	2107
	Junctional ST depression, probably normal

	2108
	Abnormal junctional ST depression

	2111
	Possible anterior injury or acute infarct

	2112
	Anterior injury or acute infarct

	2113
	Possible anteroseptal injury or acute infarct

	2114
	Anteroseptal injury or acute infarct

	2115
	Possible anterolateral subepicardial injury

	2116
	Anteroseptal subepicardial injury

	2117
	Possible septal subepicardial injury

	2118
	Septal subepicardial injury

	2119
	Possible lateral subepicardial injury

	21110
	Lateral subepicardial injury

	21111
	Possible inferio injury or acute infarct

	21112
	Inferio injury or acute infarct

	2121
	T wave abnormality, possible anterior ischemia

	2122
	T wave abnormality, possible anterior ischemia or digitalis effect

	2123
	T wave abnormality, consistent with anterior ischemia

	2124
	T wave abnormality, possible anterolateral ischemia

	2125
	T wave abnormality, possible anterolateral ischemia or digitalis effect

	2126
	T wave abnormality, consistent with anterolateral ischemia

	2127
	T wave abnormality, possible lateral ischemia

	2128
	T wave abnormality, possible lateral ischemia or digitalis effect

	2129
	T wave abnormality, consistent with lateral ischemia

	21210
	T wave abnormality, possible inferio ischemia

	21211
	T wave abnormality, possible inferio ischemia or digitalis effect

	21212
	T wave abnormality, consistent with inferio ischemia

	2131
	ST elevation, probably early repolarization

	2132
	Early repolarization

	2133
	ST elevation, consistent with subepicardial injury, pericardiatis, or Early repolarization

	2141
	Possible acute percarditis

	2142
	acute percarditis

	2151
	Nonspecific ST&T wave abnormality

	2152
	Nonspecific ST&T wave abnormality, probably digistalis effect

	2153
	Tall T waves, possible hyperkalemia

	2154
	Nonspecific T wave abnormality

	2155
	Nonspecific T wave abnormality, probably digitalis effect

	2161
	Nonspecific ST elevation

	9  Others

	911
	Low voltage

	912
	Low voltage in limb leads

	913
	Low voltage in chest leads

	941
	Long QT interval

	942
	Short QT interval

	971
	Dextrocardia ?

	972
	LIMB LEADS REVERSED

	973
	Abnormal QRS-T angle

	974
	Consistent with pulmonary disease

	981
	Artifacts present

	982
	Cannot be analyzed, re-record recommended

	10  Overall Judgment

	1010
	Normal ECG

	1011
	Borderline ECG

	1012
	Atipical ECG

	1013
	Abnormal rhythm ECG

	1014
	Abnormal ECG


· CAUTION
List of Interpretation codes may be subject to changes without notice.
Appendix E Measurement, Diagnosis, Analysis and Assessment of ECG Machine

E.1 Methods to determine the amplitude of P, QRS, ST and T wave

(1) P wave amplitude

[image: image104.emf]P

P'

P On P Off


P On is the starting position of P wave, P Off is the ending position of P wave, and dashed line is the reference baseline

To measure P wave amplitude: the line from the starting point to the ending point of P wave is the reference baseline, as shown in Figure E-1. The positive amplitude is from the reference baseline to top edge of P wave; the negative amplitude is from the reference baseline to bottom edge of P wave.

(2) QRS complex, ST segment and T wave amplitude

When measuring QRS complex, ST segment and T wave amplitude, the horizontal line of QRS complex beginning part is used as the reference baseline, as shown below:

[image: image105.emf]QRS Onset

ST deviation measured with 

respect to QRS onset

T wave amplitude measured 

with respect to QRS onset


The measurement of QRS complex, ST segment and T wave amplitude uses the horizontal line of QRS complex beginning part as the reference baseline QRS Onset is the starting position of QRS wave

E.2 Processing method of isoelectric segment in QRS complex

[image: image106.emf]P

Q

R

P

P

P'

R

R

Q

S

S

S

T

T

T

T'

Isoelectric Segment


Isoelectric segment between dash lines are in QRS complex

As shown above, the isoelectric segment beginning from the starting position of QRS complex is processed as a part of QRS complex, but doesn’t belong to the meaningful wave later (waveform area is larger than 160
[image: image107.wmf]ms

×

V

m

)

E.3 Low incidence heart disease not included in testing and diagnosis database

Test with CSE database, but this database doesn’t have sufficient number of acute myocardial infarction and myocardial ischemia ECG.

E.4 ECG diagnosis categories and the number of ECG test of each category

The accuracy of disease diagnosis and non-ECG means used to verify the effectiveness of heart disease diagnosis, as well as the patients statistics data (e.g. age, gender, race) of each group.

Test with CSE database, Table E-1 lists disease diagnostic categories, the number of ECG testing of each category and the accuracy of disease diagnosis.

CSE database sample properties are as follows:

Total number of samples: 1220 (male: 831, female: 389)

Race: White

Age: 52
[image: image108.wmf]±

13

	Type of disease
	ECG test number
	Sensitivity (%)
	Specificity (%)
	Positive predictive value (%)

	Normal
	382
	92.7
	73.9
	61.8

	Left ventricular hypertrophy
	183
	60.1
	97.0
	77.7

	Right ventricular hypertrophy
	55
	32.7
	99.9
	92.3

	Biventricular hypertrophy
	53
	26.4
	99.9
	93.3

	Anterior wall myocardial infarction
	170
	80.6
	97.7
	85.1

	Inferior wall myocardial infarction
	273
	67.0
	97.8
	89.7

	Composite myocardial infarction
	73
	64.7
	99.7
	94.0

	Hypertrophy and myocardial infarction
	31
	46.8
	100.0
	100.0


E.5 The smallest waveform identified by the device and the stability of measurement when noise exists

If the area of certain waveform is greater than or equal to 160
[image: image109.wmf]ms

×

V

m

, it is considered as meaningful wave, otherwise it is meaningless. Recognizing meaningful waveforms in area method can effectively reduce the noise. The stability of the measurement when noise exists is shown below

	Overall measurement parameter
	Type of added noise
	Mean difference (ms)
	Variance (ms)

	P time limit
	High frequency
	-0.1
	0.64

	
	Power frequency
	0.25
	1.5

	
	Low frequency
	-2.3
	3.8

	PR interval
	High frequency
	1.6
	2.4

	
	Power frequency
	-0.1
	1.5

	
	Low frequency
	0.38
	9.5

	QRS time limit
	High frequency
	0.75
	4.0

	
	Power frequency
	-1.1
	1.7

	
	Low frequency
	0.3
	4.4

	QT interval
	High frequency
	-1.6
	3.6

	
	Power frequency
	-0.5
	1.2

	
	Low frequency
	4.9
	5.6


E.6 Low incidence cardiac rhythm not included in the ECG rhythm test database

The low incidence cardiac rhythms not included in the test database:

1. Grade II conduction block;

2. Grade III conduction block.

E.7 ECG rhythm diagnosis categories and ECG test number of each category
Accuracy of rhythm diagnosis and the patient statistics data (e.g. age, gender, race) of each group

Table E-3 gives the rhythm categories, ECG test number of each category and accuracy of disease diagnosis.

The test database sample properties are as follows:

Total number of sample: 4500 (male: 2847, female: 1653)

Race: Yellow

Age: 48
[image: image110.wmf]±

12
	Rhythm type
	ECG test number
	Sensitivity (%)
	Specificity (%)
	Positive predictive value (%)

	Sinus rhythm
	3656
	98.0
	91.1
	97.9

	Premature ventricular contractions
	351
	87.2
	98.9
	81.2

	Supraventricular premature beats
	247
	68.8
	99.6
	89.9

	Atrial fibrillation
	192
	89.6
	98.7
	91.0

	Atrial flutter
	49
	65.3
	99.9
	88.9

	Pacemaker rhythm
	5
	100.0
	100.0
	100.0


E.8 Sensitivity regularly test instructions

Inspect ECGs: EGC-1C
Inspection methods:

1) Make ECG machine set in lead I, the sensitivity is set as 10 mm/mV., EGC-1C transmits the Uin as 1 mV, frequency 10 Hz sine wave signal to the ECG machine.

2) Test the waveform amplitude hm on the Inspected ECG machine. Calculate the corresponding to deviations of the sensitivity according to the following formula, should meet the maximum allowable relative deviation of ± 5 %.

[image: image111.wmf]%

100

-

n

n

m

´

=

S

S

S

S

d


The formula: Sn - nominal value of Sensitivity;

           Sm-test value of sensitivity;

           hm-the waveform amplitude of sensitivity;

           Uin-input signal amplitude if the inspected ECG machine 

3) Make ECG machine set in lead I, the sensitivity is set as 20 mm/mV. EGC-1C transmits the Uin as 0.5 mV, frequency 10 Hz sine wave signal to the ECG machine. Using the same method to test the relative deviation of 20 mm/mV sensitivity.

4) Make ECG machine set in lead I, the sensitivity is set as 5 mm/mV. EGC-1C transmits the Uin as 2 mV, frequency 10 Hz sine wave signal to the ECG machine. Using the same method to test the relative deviation of 5 mm/mV sensitivity.

5) Make ECG machine set in lead I, the sensitivity is set as 2.5mm/mV. EGC-1C transmits the Uin as 4 mV, frequency 10 Hz sine wave signal to the ECG machine. Using the same method to test the relative deviation of 2.5 mm/mV sensitivity.

6) According to the 1 and 2 steps to change the leads of the ECG machine, and make the ECG-1C's output signals connected to corresponding lead of the ECG machine, to complete all channel’s inspect, and then select the largest relative deviation from the test results for each test point, as the result of the inspection.
E.9 Distortion test

The function of ECG machine will not be affected adversely by the running of the pacemaker, which can be verified in the following way:

a) Superimpose the pulse wave of 200 mV peak, rise time less than 100 μs, 1ms pulse width, and 100 beats / min repetition rate with the sine wave signal of 1mV peak-valley value and 40 Hz frequency, and input to the ECG machine (set to standard sensitivity). The time required to restore the sine wave signals recorded by the ECG machine to 70 % of the initial value (when peak-valley value is 1mV and gain is 10 mm/mV, the initial value should be 10 mm) shouldn’t exceed 50 ms; in the above test, the maximum baseline drift accumulated in 10 s doesn’t exceed 10 mm; both with and without pulse, the amplitude difference recorded by sine wave signals (after waveform is stable) isn’t greater than ±1 mm.

b) The filter of ECG machine must be opened for distortion tests.

c) The ECG machine can pass one of the following two tests:

· String the pacemaker pulse wave of 200 mV peak, rise time less than 100 μs, 1ms pulse width, and 120 pulses / min repetition rate together with the symmetrical triangular wave of 2 mV amplitude and 100 ms duration. The starting time of pulse wave should be 40 ms earlier (or later) than the starting time of triangular wave, input such a signal to the ECG machine, record in the standard sensitivity, the triangular wave is clearly visible on the ECG machine records, the difference between recorded amplitude and the original amplitude (the original amplitude of the waveform with 2 mV amplitude should be 20 mm under 10 mm/mV gain) does not exceed 20 %, and the location of the pacemaker pulse can be clearly identified in the ECG machine records.

· String the pacemaker pulse wave of 200 mV peak, rise time less than 100 μs, 1ms pulse width, and 120 pulses / min repetition rate together with the ECG calibration signal CAL20000, and input to the ECG machine. The QRS curve of calibration signal can be clearly identified on ECG machine records, the difference between the recorded amplitude and the original amplitude of QRS curve does not exceed 20%, and the location of the pacemaker pulse can be clearly identified in the ECG machine records

Appendix F EMC-Guidance and manufacture’s declaration

· CAUTION

· The use of accessories, sensor and cable which exceed regulations can increase the electromagnetic emissions or decrease the electromagnetic immunity.

· Please don’t use the equipment with tightness and pile. The equipment should be detected according to the requirements to make sure it working normally.

· Please make the equipment safety according to the EMC, and install, fix it under the situation required by EMC.

· Even though the other equipments which measure up the CISPR also may cause disturbance.

· When the amplitude of import signal narrow than regulated minimum amplitude may lead to inaccurate measure.
· Portable and mobile communications equipment will affect the performance of the equipment.
F.1 Guidance and manufacturer’s declaration – electromagnetic emission

	Guidance and manufacturer’s declaration – electromagnetic emission

	The Digital Electrocardiograph is intended for use in the electromagnetic environment specified below. The customer or the user of the ECG machine should assure that it is used in such an environment.

	Emissions test
	Compliance
	Electromagnetic environment - guidance

	RF emissions


EN 55011
	Group 1
	The ECG machine uses RF energy only for its internal function. Therefore, its RF emissions are very low and are not likely to cause any interference in nearby electronic equipment.

	RF emissions

EN 55011
	3-Channel Machine: Class B
	The ECG machine is suitable for use in all 

establishments, including domestic establishments and those directlyconnected to the public low-voltage power supplynetwork that 

supplies buildings used for domestic purposes.

	
	6-Channel Machine: Class A
	

	Harmonic emissions
EN 61000-3-2
	Class A
	

	Voltage fluctuations /
flicker emissions
EN 61000-3-3
	Complies
	


F.2 Guidance and manufacture’s declaration-electromagnetic immunity for all EQUIPMENT and SYSTEMS.

	Guidance and manufacture’s declaration-electromagnetic immunity

	The ECG machine is intended for use in the electromagnetic environment specified below. The customer of the user of the ECG machine should assure that it is used in such an environment

	Immunity test
	IEC60601 test level
	Compliance level
	Electromagnetic environment guidance

	Electrostatic discharge(ESD)

IEC61000-4-2
	±6KV contact

±8KV air
	±6KV contact

±8KV air
	Floors should be wood, concrete or ceramic tile. If floor are covered with synthetic material, the relative humidity should be at least 30%

	Electrical fast transient/burst IEC61000-4-4  
	±2KV for power supply lines

±1KV for input/output lines
	±1KV for power supply  lines

±1KV for input/output lines
	Mains power quality should be that of a typical commercial or hospital environment.

	Surge IEC61000-4-5
	±1KV differential mode

±2KV common mode
	±1KV differential mode

±2KV common mode
	Mains power quality should be that of a typical commercial or hospital environment.

	Voltage dips, short interruptions and voltage variation power supply input lines IEC61000-4-11
	<5% UT

(>95% dip in UT)

for 0.5 cycle

40% UT

(60% dip in UT)

for 5 cycle

70% UT

(30% dip in UT)

for 25 cycles

<5% UT

(>95% dip in UT) 

for 5 sec  
	<5% UT

for 0.5 cycle

40% UT

for 5 cycle

70% UT

for 25 cycles

<5% UT

for 5 sec  
	Main power quality should be that of a typical commercial or hospital environment. If the user of the ECG machine requires 

Continued operation during power mains interruptions, it is recommended that the ECG machine be powered from an uninterruptible power supply or a battery.

	Power frequency

(50Hz)magnetic

field

IEC61000-4-8
	3A/m
	3A/m
	Mains power quality should be that of a typical commercial or hospital environment.

	NOTE： UT is the A.C mains voltage prior to application of the test level.


F.3 Guidance and manufacture’s declaration-electromagnetic immunity for EQUIPMENT and SYSTEMS those are not LIFE-SUPPORTING

	Guidance and manufacture’s declaration-electromagnetic immunity

	The ECG machine is intended for use in the electromagnetic environment specified below. The customer of the user of the ECG machine should assure that it is used in such an environment.

	Immunity test
	IEC60601 test level
	Compliance level
	Electromagnetic environment-guidance

	Conducted RF

IEC61000-4-6

Radiated RF

IEC61000-4-3
	3 Vrms

150kHz -80 MHz

3 V/m

80 MHz -2.5 GHz


	3 Vrms

3 V/m


	Portable and mobile RF communications equipment should be used no closer to any part of the ECG machine, including cables, than the recommended separation distance calculated from the equation applicable to the frequency of the transmitter.

Recommended separation distance

[image: image112.wmf]P

V

5

.

3

d

1

ú

û

ù

ê

ë

é

=


[image: image113.wmf]P

E

5

.

3

d

1

ú

û

ù

ê

ë

é

=

80MHz-800 MHz

[image: image114.wmf]P

E

7

d

1

ú

û

ù

ê

ë

é

=

800MHz-2.5 GHz

Where P is the maximum output power rating of 

the transmitter in watts (W) according to the 

transmitter manufacturer and d is the 

recommended separation distance in metres (m).

Field strengths from fixed RF transmitters, as

determined by an electromagnetic site survey, 

range a  should be less than the compliance level

 in each frequency range.b

Interference may occur in the vicinity of 

Equipment marked with the 

following symbol:

[image: image115.png](t
{ g(o)))





	NOTE1：At 80MHz and 800MHz, the higher frequency range applies.

NOTE2：These guidelines may not apply in all situations. Electromagnetic propagation is affected by absorption and reflection from structures, objects and people.

	A．Field strengths from fixed transmitters, such as base stations for radio(cellular/cordless) telephones and land mobile radios, amateur radio, AM and FM radio broadcast and TV broadcast cannot be predicted theoretically with accuracy. To assess the electromagnetic environment due to fixed RF transmitters, an electromagnetic site survey should be considered. If the measured field strength in the location in which the ECG machine should be observed to verify normal operation. If abnormal performance is observed, additional measures may be necessary, such as reorienting or relocating the ECG machine.

B．Over the frequency range 150kHz to 80 MHz, field strengths should be less than 3V/m.


F.4 Recommended separation distance between portable and mobile RF communications equipment and the EQUIPMENT or SYSTEM for EQUIPMENT or SYSTEM that are not LIFE-SUPPORTING

	Recommended separation distance between portable and mobile RF communications equipment and the ECG machine

	The ECG machine is intended for use in an electromagnetic environment in which radiated RF disturbances are controlled. The customer or the user of the ECG machine can help prevent electromagnetic Interference by maintaining a minimum distance between portable and mobile RF communications equipment (transmitters) and the ECG machine as recommended below according to the maximum output power of the communications equipment.

	Rated maximum output

power of transmitter

(W)
	Separation distance according to frequency of transmitter(m)

	
	150 kHz-80 MHz 

[image: image116.wmf]P

V

5

.

3

d

1

ú

û

ù

ê

ë

é

=


	80 MHz-800 MHz

[image: image117.wmf]P

E

d

ú

û

ù

ê

ë

é

=

1

5

.

3


	800 MHz-2.5 GHz

[image: image118.wmf]P

E

7

d

1

ú

û

ù

ê

ë

é

=



	0.01
	0.12
	0.12
	0.23

	0.1
	0.38
	0.38
	0.73

	1
	1.2
	1.2
	2.3

	10
	3.8
	3.8
	7.3

	100
	12
	12
	23

	For transmitters rated at a maximum output power not listed above, the recommended separation distance in meters (m) can be estimated using the equation applicable to the frequency of the transmitter, where P is the maximum output power rating of the transmitter in watts (W) according to the transmitter manufacturer.

NOTE1： At 80MHz and 800MHz, the higher frequency range applies.

NOTE2： These guidelines may not apply in all situations. Electromagnetic propagation is affected by absorption and reflection from structures, objects and people.


[image: image119.jpg]


[image: image120.jpg]


[image: image121.jpg]&



[image: image122.png]mc




®





Acquisition Module





Lead Wires





Card Slot





ECG machine





Switchboard





Server





Wireless AP





Server





Switchboard





ECG machine





Unipolar Chest leads





Unipolar Limb leads





Standard Limb leads





�





Entering patient information 





�





Set the recording mode and recording format





Observe the waveforms displayed





Abnormal(Lead off, data overflow and etc)








Normal





Refer to Troubleshooting to find solutions.





[Print]/�





Recording an ECG





[Stop]/�





Exit the recording





Recording Waveforms and report





Arrhythmia is detected.   


Print the rhythm waveforms?





[Yes]





[No]





Return to main interface





Print rhythm waveforms





Recording


waveforms





Manual Report





Simple Report/Detail Report/Median Beat Report


代表心搏报告





Disable





Disable


Simple Report Detail Report


Median Beat Report Manual Report








Edit, and then print





Print selected report





No report printed


报告





Main interface





[Rhythm]


】





Enter rhythm interface





Enter rhythm report interface





Main interface





[Report]





Enter analysis report





Beat waveforms





Minnesota code





Simple measurement data





Analysis result: press here, you can change the analysis result.





Rhythm waveforms





Main interface





[Freeze]





Enter freeze interface





File Attributes Area





Indication Area





ECG Files


Display Area





Select All


Shortcut Button





ECG files are copied or moved to the selected medium





Choose destination medium





Select [Copy] or [Move]





[Operate]





Select one or several files





[Data Management]








_1476606617.unknown

_1480919050.bmp




_1483772421.bmp




_1492499333.vsd
�

�

�

�

�

�

�

�

�

�

�

￼�

DC-DC converter


Switch power source


Charger


DC-DC converter


CF-type part


Logic switch


Recorder


DC-DC converter


Keyboard


Reset


Control system


Battery
14.8V


AC 100-240V
50/60Hz


±8V


+5V


+24V


+5V


-5V


+3.3V/+1.8V


(+12V)



_1483542216.bmp
Card_Slot




_1476606619.unknown

_1476606620.unknown

_1476606618.unknown

_1476606615.vsd
Main Unit


DC/DC
Isolated power


Signal isolation


Microprocessor


Acquisition module keys


Lead wire


Analog circuit section



_1476606616.vsd
Recorder


Program memory


Analog board


LCD


LCD controller


Programmable logic unit


Microprocessor


Interface


CRO


EXT


Real time clock


Data memory


Keyboard



_1473168659.png




